ALMA Front End

Design Specifications and Requirements

Doc. No.: FEND-40.00.00.00-001-A-SPE

Version 2.5 (Draft)

2003-05-28

Prepared:	Hans Rudolf, ALMA Front End IPT

ESO				2003-05-28

Name				Date				Signature

Approved:

Name				Date				Signature

Released:

Name				Date				Signature

Revision history

| Issue | Approval Date | Part(s) affected | Reason/remarks |
| --- | --- | --- | --- |
| 1 |  | All | First document release |

### TABLE OF CONTENTS

1.	Introduction	7

1.1.	Purpose	7

1.2.	Scope	7

1.3.	Applicable documents	9

1.4.	Reference documents	9

1.5.	Acronyms	10

2.	Description	11

2.1.	Equipment Definition	11

2.2.	Definition of terms	11

2.3.	Block diagram	13

3.	Functional Requirements	15

3.1.	Operation modes	15

3.1.1.	Operational	15

3.1.2.	Non-Operational	15

3.1.3.	Stand-by	15

3.1.4.	Transport with the antenna	15

3.1.5.	Transport in the Front End service vehicle	15

3.1.6.	Storage	15

3.2.	General	15

3.2.1.	Pre-selection of observation bands	15

3.2.2.	Mechanical tuning	16

3.2.3.	Standard parts	16

3.2.4.	Cables and connectors	16

3.2.5.	Solar observing and safety	16

3.3.	Frequency Coverage	16

3.3.1.	RF input port	16

3.3.2.	LO input port	17

3.3.3.	IF output port bandwidth and centre frequency	17

4.	Performance requirements	18

4.1.	Cryogenics, Dewar, and Vacuum	18

4.1.1.	Evacuation and cool-down time	18

4.1.2.	Warm-up time	18

4.1.3.	Re-pump period	18

4.2.	Optics	18

4.2.1.	Polarisation	18

4.2.1.1.	Polarisation Optimisation	18

4.2.1.2.	Polarisation States	18

4.2.1.3.	Polarisation alignment accuracy	18

4.2.1.4.	Cross-Polarisation	18

4.2.1.5.	Polarisation mismatch	18

4.2.2.	Beam pattern / Beam efficiency	19

4.2.3.	Insertion Loss	19

4.2.4.	Pointing stability	19

4.2.5.	Solar Filter	19

4.2.6.	Quarter-Wave-Plate	19

4.2.7.	Amplitude Calibration	19

4.3.	Water Vapour Radiometer	19

4.3.1.	WVR beam accuracy	19

4.3.2.	WVR sensitivity	19

4.3.3.	WVR stability	20

4.3.4.	WVR tuning range and step size	20

4.3.5.	WVR calibration interval	20

4.4.	Front End Noise Performance	20

4.5.	Image Band Suppression / sideband mismatch	21

4.6.	Spurious response of the receiver	21

4.7.	Out-of band response	21

4.8.	Saturation	21

4.9.	Gain	21

4.10.	Gain stability	22

4.11.	Signal path phase stability	22

5.	Mechanical and electrical Requirements	23

5.1.	Mass	23

5.2.	Centre of Gravity	23

5.3.	Eigen-frequencies	23

5.4.	Volume	23

5.5.	Orientation	23

5.6.	Thermal Load	23

5.7.	Power requirements	23

6.	operating conditions	24

6.1.	Stabilisation time	24

6.2.	Repeatability	24

6.3.	Simultaneous operation of bands	24

6.4.	Band Selection	24

6.4.1.	Selection of a (pre-set) observing band	24

6.4.2.	Selection of new observing band	24

6.4.3.	Narrow-band frequency switching	24

6.4.4.	Frequency changes within a band	24

6.5.	Local Oscillator	25

6.5.1.	LO phase stability	25

6.5.2.	LO phase resolution	25

6.5.3.	LO Phase switch step size	25

6.5.4.	LO Phase switching time	25

6.6.	Monitoring and Control	25

6.7.	Environmental operating conditions	25

6.7.1.	Altitude	25

6.7.2.	Thermal Environment	25

6.7.3.	Relative Humidity	25

6.7.4.	Vibration	25

6.7.5.	Acceleration-	26

6.7.6.	Cleanliness	26

6.8.	Storage and shipping conditions	26

6.9.	EMC	26

6.9.1.	RFI	26

6.10.	Grounding / Isolation	26

6.11.	Availability, reliability and maintainability	26

6.11.1.	MTBF	26

6.11.2.	MTTR	26

6.11.3.	Lifetime	26

6.11.4.	Maintenance	27

6.11.5.	Hold times	27

# Introduction

## Purpose

This document summarizes the key design specifications and requirements for the Front End. The Front End is a cryogenically cooled 10 band receiver that converts radio frequencies ranging from 31.3 GHz to 950 GHz to intermediate frequencies from 4 to 12 GHz. It includes auxiliary radio beam components, a cryogenic cooling sub-system, a water vapour radiometer system as well as monitoring and control electronics.

This specification has been prepared for the detailed design of the ALMA Front End in ALMA phase 2.

## Scope

The information given in this document provides a complete summary of all the requirements and specifications that must be met by the Front End design. The specifications are presented in a “black box” view. This document is accompanied by the specifications of the Front End units and Interface Control Documents for the Front End.

The following table shows a partial view of the ALMA product tree [1] at “module” and “unit” level for the ALMA Front End products that are covered by this document. Those products belonging to the FE sub-system that are not specified by this document are clearly identified in the following table.

| PT level 1 / “sub-system” |  | PT level 2 / “module” |  | PT level 3 / “unit” |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Product No. | Product Name | Product No. | Product Name | Product No. | Product Name | Remarks |
| 40.00.00.00 | Front end |  |  |  |  |  |
|  |  | 40.01.00.00 | Warm optics |  |  |  |
|  |  | 40.02.00.00 | Cartridges |  |  |  |
|  |  |  |  | 40.02.01.00 | Frequency band 1 cartridge |  |
|  |  |  |  | 40.02.02.00 | Frequency band 2 cartridge |  |
|  |  |  |  | 40.02.03.00 | Frequency band 3 cartridge |  |
|  |  |  |  | 40.02.04.00 | Frequency band 4 cartridge |  |
|  |  |  |  | 40.02.05.00 | Frequency band 5 cartridge |  |
|  |  |  |  | 40.02.06.00 | Frequency band 6 cartridge |  |
|  |  |  |  | 40.02.07.00 | Frequency band 7 cartridge |  |
|  |  |  |  | 40.02.08.00 | Frequency band 8 cartridge |  |
|  |  |  |  | 40.02.09.00 | Frequency band 9 cartridge |  |
|  |  |  |  | 40.02.10.00 | Frequency band 10 cartridge |  |
|  |  | 40.03.00.00 | Cryostat |  |  |  |
|  |  |  |  | 40.03.01.00 | Dewar |  |
|  |  |  |  | 40.03.02.00 | Cryo-cooler |  |
|  |  |  |  | 40.03.03.00 | Vacuum pumps |  |
|  |  |  |  | 40.03.04.00 | Cryostat electrical infrastructure |  |
|  |  | 40.04.00.00 | Front end auxiliary sub-systems |  |  |  |
|  |  |  |  | 40.04.01.00 | Front end power supply sub-system |  |
|  |  |  |  | 40.04.02.00 | Bias electronics sub-system |  |
|  |  |  |  | 40.04.03.00 | Front end M&C sub-system |  |
|  |  | 40.05.00.00 | Front end chassis |  |  |  |
| PT level 1 / “sub-system” |  | PT level 2 / “module” |  | PT level 3 / “unit” |  |  |
| Product No. | Product Name | Product No. | Product Name | Product No. | Product Name | Product No. |
|  |  |  |  | 40.05.01.00 | Front end mechanical structure |  |
|  |  |  |  | 40.05.02.00 | Front end cabling |  |
|  |  | 40.06.00.00 | Front end integrated calibration & widgets |  |  |  |
|  |  |  |  | 40.06.01.00 | Vane calibration sub-system |  |
|  |  |  |  | 40.06.02.00 | Solar protection |  |
|  |  |  |  | 40.06.03.00 | Polarisation widgets |  |
|  |  | 40.07.00.00 | Water vapour radiometer |  |  |  |
|  |  | 40.08.00.00 | Front end IF |  |  |  |
|  |  |  |  | 40.08.01.00 | IF switch sub-system |  |
|  |  | 40.09.00.00 | Front end specific test, construction & service equipment |  |  | Not covered in this document |
|  |  |  |  | 40.09.01.00 | SIS mixer fabrication equipment | Not covered in this document |
|  |  |  |  | 40.09.02.00 | SIS mixer test equipment | Not covered in this document |
|  |  |  |  | 40.09.03.00 | Front end test fixture | Not covered in this document |
|  |  |  |  | 40.09.04.00 | Cartridge test dewars | Not covered in this document |
|  |  |  |  | 40.09.05.00 | Cartridge RF test fixtures | Not covered in this document |
|  |  |  |  | 40.09.06.00 | Front end service vehicle | Not covered in this document |
|  |  | 40.10.00.00 | First local oscillator |  |  |  |
|  |  |  |  | 40.10.01.00 | First LO frequency sources |  |
|  |  |  |  | 40.10.02.00 | Warm frequency multipliers |  |
|  |  |  |  | 40.10.03.00 | First LO PLL unit |  |
|  |  |  |  | 40.10.04.00 | Band selection |  |

## Applicable documents

The following documents are included as part of this document to the extent specified herein. If not explicitly stated differently, the latest issue of the document is valid.

| Reference | Document title |  |  | Document ID |
| --- | --- | --- | --- | --- |
| [AD1] | ALMA Product Tree |  |  | SYSE-80.03.00.00-001H-LIS http://almaedm.tuc.nrao.edu/forums/alma/dispatch.cgi/approjsys/docProfile/100005/d20030325213221/No/t100005.htm |
| [AD2] | ALMA Equipment Shipping & Storage Environmental Specification |  |  | ALMA-80.05.02.00-004-A-SPE http://almaedm.tuc.nrao.edu/forums/alma/dispatch.cgi/docsyseng/docProfile/100351/d20021118154650/No/t100351.htm |
| [AD3] | ALMA system wide requirements – Electromagnetic Compatibility |  |  | ALM-SPE-ESO-12000-00007 http://almaedm.tuc.nrao.edu/forums/alma/dispatch.cgi/docsyseng/docProfile/100439/d20030207161546/No/t100439.htm |
| [AD4] | ICD | Antenna | Front End | (in preparation) |
| [AD5] | ICD | Front End | Back End | (in preparation) |
| [AD6] | ICD | Front End | Computing | (in preparation) |
| [AD6] | ICD | Site | Front End | (in preparation) |
| [AD7] | ALMA Electronic Design Specifications and Guidelines |  |  | ALMA-80.05.00.00-005-A-SPE http://almaedm.tuc.nrao.edu/forums/alma/dispatch.cgi/docsyseng/docProfile/100405/d20030125001225/No/t100405.htm |
| [AD8] | ALMA Power Quality Specification |  |  | SYSE-80.05.00.00-001A-SPE http://almaedm.tuc.nrao.edu/forums/alma/dispatch.cgi/docsyseng/docProfile/100423/d20030206161457/No/t100423.htm |
| [AD9] | ALMA Lifetime Definitions of ALMA Products |  |  | SYSE-80.05.00.00-002A-SPE http://almaedm.tuc.nrao.edu/forums/alma/dispatch.cgi/docsyseng/docProfile/100432/d20030206163550/No/t100432.htm |

In the event of a conflict between one of the above referenced applicable documents and the contents of this document, the contents of this document shall be considered as a superseding requirement.

## Reference documents

The following documents contain additional information and are referenced in this document.

| Reference | Document title | Document ID |
| --- | --- | --- |
| [1] | List of acronyms and glossary for the ALMA project | (in preparation) |
| [2] | ALMA Project Book | http://www.alma.nrao.edu/projectbk/construction/ |
| [3] | ALMA Receiver Optics Design | ALMA Memo #362 http://www.alma.nrao.edu/memos/html-memos/abstracts/abs362.html |

## Acronyms

A limited set of basic acronyms used in this document is given below. A complete set of acronyms used in the ALMA project can be found in reference [1].

ALMA	Atacama Large Millimetre Array

CDR	Critical Design Review

DSB	Double-SideBand

FESS	Front End Support Structure

#### ICD		Interface Control Document

LO		Local Oscillator

MTBF	Mean Time Between Failures

MTTR	Mean Time To Repair

PDR	Preliminary Design Review

SSB	Single-SideBand

WVR	Water Vapour Radiometer

2SB	Dual-SideBand separating

# Description

## Equipment Definition

The Front End (FE) assembly accepts the focused beam from the antenna's secondary reflector over a selected band of frequencies.  It amplifies and converts this band to an intermediate frequency band in several channels (typically differing in polarisation and/or sideband), and delivers the IF signals as outputs.  It accepts as inputs local oscillator reference signals at the appropriate frequencies and levels to generate the local oscillator signal for conversion.  The Front End assembly includes:

- RF optics as required to couple the sub-reflector beam to its first electronic element.

- Mixers and amplifiers of the RF-IF signal path (separately for each band required to cover the overall frequency range), including any components needed to couple the LO signal to each mixer.

- Local oscillator components within the cryostat and associated hardware. The interface between the LO subsystem and the Front End subsystem is located at the waveguide input to the warm multiplier assembly at frequencies of around 100 GHz. Each cartridge will have its own interface to the LO sub-system. The interface will be specified in detail in a corresponding ICD.

- Vacuum system and cryo-coolers needed to achieve the appropriate operating temperatures for certain components, along with related thermal insulation and mechanical supports.

- Bias and control circuits to support the proper operation of the RF-IF amplifiers and mixers.

- Devices to select the desired frequency, including IF band switching and any required electronic configuration of amplifiers and mixers as a function of frequency.

- A local monitor and control system permitting remote control of all functions and providing extensive remote diagnosis capability, with an appropriate interface to the general ALMA Monitor and Control bus. Design implementations of such a system will be decided by ICDs between effected systems.

- A water-vapour monitoring radiometer.

- Devices that are placed directly in the input radio beam of the receiver and which include (but are not limited to) calibration systems, insertable components such as quarter wave plates for the reception of circular polarisation, and attenuators for solar observations.

It does not include the following elements, which belong to other subsystems:

- the local oscillator sub-system up to the specified interface (see ICD) and

- calibration devices located outside the receiver cabin (including any built into the sub-reflector).

This specification also does not cover the following elements, though they are a Front End IPT responsibility:

- the service and exchange vehicle which are used to transport and facilitate the installation of the Front End and

- Front End specific test and construction equipment.

## Definition of terms

- Front end assembly: this is the Front End subsystem as defined in 2.1. It provides space for 10 cartridges (see below).

- Band: a range of RF signal frequencies that is received in orthogonal linear polarisation and defined in 3.3.1.

- Frequency channel: this is one receiving chain that receives one polarisation within the specified band. A band has two channels.

- Cartridge: a device insertable into the main dewar, receiving RF frequencies within one specified band in dual polarisation, containing optics, mixers, IF amplifiers, LO components. Receives the RF from external optics, accepts an LO signal, and delivers IF signals.

- ICD: Interface Control Document. The document that specifies interfaces between sub-systems and parts within a sub-system.

## Block diagram

# Functional Requirements

## Operation modes

The ALMA Front End can exist in the following states.

## Operational

FEND-91010-ZZZ

This mode defines the state in which operational power is applied with all active signal levels at nominal values. All specifications and requirements in this document apply, unless otherwise stated. The operational mode is also called “On”.

## Non-Operational

FEND-91020-ZZZ

This mode defines the state in which operational power is not applied and signal levels are not at nominal values. The operational mode is also called “Off”.

## Stand-by

FEND-91025-ZZZ

This mode defines the state in which operational power is applied to a cartridge within the Front End and cartridge signal levels are not at nominal values. The stand-by mode applies only to cartridges within the Front End and not to the whole Front End assembly.

## Transport with the antenna

FEND-91030-ZZZ

The mode defines the state in which the Front End is transported with the antenna on the antenna transport vehicle. This mode differs from the Non-operational mode basically in the environmental operating conditions. For this mode, the same specifications and requirements as for the non-operational mode apply, unless otherwise stated.

## Transport in the Front End service vehicle

FEND-91040-ZZZ

The mode defines the state in which the Front End is transported with the Front End service and exchange vehicle.  This mode differs from the Non-operational mode basically in the environmental operating conditions.  For this mode, the same specifications and requirements as for the non-operational mode apply, unless otherwise stated.

## Storage

FEND-91050-ZZZ

The mode defines the state in which the ALMA Front End is stored completely assembled. This mode differs from the Non-operational mode basically in the environmental conditions and the lack of monitoring and control signals. For storage, the same specifications and requirements as for the non-operational mode apply, unless otherwise stated.

## General

## Pre-selection of observation bands

FEND-92110-ZZZ

Means shall be provided for the application of nominal power and signals to the desired cartridge.  Specifications and requirements of this document no do fully apply.

## Mechanical tuning

FEND-92020-ZZZ

No mechanical tuning shall be employed during operation.

## Standard parts

FEND-92030-ZZZ

Standard, unmodified commercially available components shall be used where possible.

## Cables and connectors

FEND-92040-ZZZ

All cables and connectors shall be in compliance with [AD7].

## Solar observing and safety

FEND-62910-ZZZ

No components shall be damaged with 0.3 W/cm2 of solar radiation incident anywhere on a Front End assembly.  This specification applies for all modes.

Provisions shall be taken to allow observations of the sun.

## Frequency Coverage

This section applies only to the operational mode.

## RF input port

| Band | Start frequency | Stop frequency | Remarks |  |
| --- | --- | --- | --- | --- |
| 1 | 31.3 GHz | 45 GHz |  | FEND-21310-ZZZ |
| 2 | 67 GHz | 90 GHz |  | FEND-22310-ZZZ |
| 3 | 86 GHz | 116 GHz | operation to 84 GHz | FEND-23310-ZZZ |
| 4 | 125 GHz | 163 GHz |  | FEND-24310-ZZZ |
| 5 | 163 GHz | 211 GHz |  | FEND-25310-ZZZ |
| 6 | 211 GHz | 275 GHz |  | FEND-26310-ZZZ |
| 7 | 275 GHz | 370 GHz | operation to 372 GHz | FEND-27310-ZZZ |
| 8 | 385 GHz | 500 GHz |  | FEND-28310-ZZZ |
| 9 | 602 GHz | 720 GHz |  | FEND-29310-ZZZ |
| 10 | 787 GHz | 950 GHz |  | FEND-20310-ZZZ |

Full specifications and requirements do not apply for operation outside nominal frequency limits.

## LO input port

| Band | Bottom frequency | Top frequency | Remarks |  |
| --- | --- | --- | --- | --- |
| 1 | 27.3 GHz | 33 GHz | USB mixing scheme | FEND-21320-ZZZ |
| 2 | 79 GHz | 94 GHz | LSB mixing scheme | FEND-22320-ZZZ |
| 3 | 94 GHz | 108 GHz | operation to 92 GHz | FEND-23320-ZZZ |
| 4 | 133 GHz | 155 GHz |  | FEND-24320-ZZZ |
| 5 | 171 GHz | 203 GHz |  | FEND-25320-ZZZ |
| 6 | 219 GHz | 267 GHz |  | FEND-26320-ZZZ |
| 7 | 283 GHz | 362 GHz | operation to 364 GHz | FEND-27320-ZZZ |
| 8 | 393 GHz | 492 GHz |  | FEND-28320-ZZZ |
| 9 | 610 GHz | 712 GHz |  | FEND-29320-ZZZ |
| 10 | 795 GHz | 942 GHz |  | FEND-20320-ZZZ |

Full specifications and requirements do not apply for operation outside nominal frequency limits.

## IF output port bandwidth and centre frequency

FEND-80310-ZZZ

Each signal channel shall provide 8 GHz total IF bandwidth per polarisation using one of the following alternatives depending on the mixing scheme selected:

  - 8 GHz bandwidth single-sideband (SSB), upper or lower sideband centred at 8.0 GHz

  - 8 GHz bandwidth double-sideband (DSB), centred at 8.0 GHz

  - 4 GHz bandwidth dual-sideband, (2SB) upper and lower sideband, centred at 6.0 GHz

# Performance requirements

## Cryogenics, Dewar, and Vacuum

## Evacuation and cool-down time

FEND-32610-ZZZ

The dewar with all cartridges installed shall be ready for operation in 48 hours (TBC) maximum, with the use of an external backing pump.

## Warm-up time

FEND-32620-ZZZ

The warm-up of the dewar to ambient temperature shall be performed in 6 hours maximum.

## Re-pump period

FEND-33610-ZZZ

Periodic dewar re-pumping can be performed at intervals of not less than 12 months.

## Optics

This section applies only to the operational mode.

## Polarisation

## Polarisation Optimisation

FEND-11010-ZZZ

The polarisation performance shall be optimised for band 7.

## Polarisation States

FEND-11110-ZZZ

The Front End shall simultaneously receive two orthogonal polarisations, with each converted to one or more separate IF outputs depending on mixing scheme. The nominal polarisation states shall be linear.

## Polarisation alignment accuracy

FEND-11210-ZZZ

The orientation of each beam’s E-vector shall not deviate more than 2o peak-to-peak from the nominal.

## Cross-Polarisation

FEND-11310-ZZZ

At any frequency within the Front End's tuning range, the cross-polarised contribution within a signal channel shall be at least 20 dB below the desired polarisation.

## Polarisation mismatch

FEND-11410-ZZZ

The Front End contribution to the maximum polarisation mismatch between any pair of antennas in the array shall not exceed -20 dB.

## Beam pattern / Beam efficiency

FEND-12010-ZZZ

The beam efficiency for all Front End Assembly bands and orthogonal polarisations shall exceed 90 % (TBC).  The Front End Assembly optical system beam efficiency is defined as the ratio of power contained in the main beam to the integrated beam power.

## Insertion Loss

FEND-13010-ZZZ

The optical coupling of each Front End channel to the telescope shall be maximized. Details can be found in [3].

## Pointing stability

FEND-14010-ZZZ

Tipping of the Front End assembly shall result in a telescope beam change of less than 0.1 arc-seconds.

Tipping of the Front End assembly shall result in a telescope coupling change of less than 2 %.

Tipping of the Front End assembly shall result in a path length change of less than 4.5 m.

## Solar Filter

FEND-62920-ZZZ

A solar filter shall be provided for solar observations. Remote controlled automatic insertion into the beam shall be provided. It shall provide at least 20 dB attenuation of 10 micron radiation and an RF signal attenuation of 13 dB for all ALMA frequency bands.

## Quarter-Wave-Plate

FEND-64010-ZZZ

An insertable quarter wave plate shall be provided for band 7. The centre frequency of the quarter wave plate shall be 345 GHz. The combined absorptive and reflective losses shall be less than 0.5 dB (TBC). The ellipticity shall be better than 10 % (TBC). These specifications shall apply over a 10 % (TBC) fractional bandwidth.

## Amplitude Calibration

FEND-65010-ZZZ

Means for periodic amplitude calibration shall be provided. Calibration shall be completed in a time not to exceed 2 s. The amplitude calibration error shall be no greater than 1 % (TBC).

## Water Vapour Radiometer

FEND-70910-ZZZ

This subsection applies only to the operational mode.

The Front End shall include a radiometer for measurement of water vapour along the signal path, using the 183 GHz line.  This instrument shall operate simultaneously with the selected astronomy band illuminating the sub-reflector. It must include all necessary LO sources, signal processing, and monitoring and control.

## WVR beam accuracy

FEND-70920-ZZZ

The beam position of the WVR relative to the observing band shall be within 10 arc-minutes.

## WVR sensitivity

FEND-70930-ZZZ

The computed rms path error of the WVR shall be less than 10*(1+wv) µm of path, with wv being the water vapour along the line of sight in millimetres. This sensitivity shall be achieved with a time resolution of 1 second.

## WVR stability

FEND-70315-ZZZ

The sensitivity of the WVR shall be maintained over time periods of 5 minutes and for changes in zenith angle of up to 1 degree.

## WVR tuning range and step size

FEND-70310-ZZZ

The tuning range of the WVR shall be no less than 25 MHz (TBC), allowing for tuning steps of no bigger than 25 kHz (TBC).

## WVR calibration interval

FEND-70610-ZZZ

The calibration interval of the WVR shall be 5 minutes.

## Front End Noise Performance

This section applies only to the operational mode.

The following table shows the required noise temperature performance of the ALMA Front End. The noise performance is referred to its effective RF input port including all contributions from warm optics, dewar windows, and IR filters. It must take into account all noise contributions up to the Front End assembly IF output ports.

Depending on the selected mixer scheme the cartridge noise temperature shall not exceed the values of either TSSB for SSB and 2SB response or TDSB = 0.5.TSSB for DSB responses as follows:

Table:	Specifications for maximum receiver noise temperatures

| Band | SSB |  | DSB |  |  |
| --- | --- | --- | --- | --- | --- |
|  | T(SSB) over 80% of the RF band | T(SSB) at any RF frequency | T(DSB) over 80% of the RF band | T(DSB) at any RF frequency |  |
| 1 | 15 K | 23 K | 8 K | 12 K | FEND-21910-ZZZ |
| 2 | 28 K | 43 K | 14 K | 22 K | FEND-22910-ZZZ |
| 3 | 34 K | 54 K | 17 K | 27 K | FEND-23910-ZZZ |
| 4 | 47 K | 76 K | 24 K | 38 K | FEND-24910-ZZZ |
| 5 | 60 K | 98 K | 30 K | 49 K | FEND-25910-ZZZ |
| 6 | 77 K | 126 K | 39 K | 63 K | FEND-26910-ZZZ |
| 7 | 133 K | 198 K | 67 K | 99 K | FEND-27910-ZZZ |
| 8 | 181 K | 270 K | 91 K | 135 K | FEND-28910-ZZZ |
| 9 | 335 K | 500 K | 168 K | 250 K | FEND-29910-ZZZ |
| 10 | 438 K | 655 K | 219 K | 328 K | FEND-20910-ZZZ |

Remarks:

- The frequency ranges of the bands in the table above are specified in section 3.3.1 of this document.

- The noise temperatures shall be achieved for the full IF band, as defined in section 3.3.3

- The noise temperature shall be calculated from measurements according to the Rayleigh-Jeans law.

- The noise temperatures in this table are based on the approved “Specifications for the ALMA Front End Assembly”, version 1.0, and the following.

- Following an ASAC recommendation the values in the Table were calculated with the following formula:

T(SSB)= A * (h*freq/k) + 4 K

where h and k are the usual physical constants, and freq was taken as the centre frequency of a particular band. The frequency dependent quantity A has the following specification and values (over 80% of the RF band / at any freq):

Bands 1-6 (below 275 GHz)  	Spec: A = 6 / 10

Bands 7-8 (275-500 GHz)    		Spec: A = 8 / 12

Band  9   (602-720 GHz)    		Spec: A = 10 / 15

Band  10  (787-950 GHz)   		Spec: A = 10 / 15

## Image Band Suppression / sideband mismatch

FEND-80910-ZZZ

This section applies only to the operational mode.

For a SSB or 2SB mixing scheme the image band suppression shall be at least 10 dB.

For a DSB mixing scheme the side-band ratio shall be less than 1.5 dB (TBC).

## Spurious response of the receiver

FEND-80920-ZZZ

This section applies only to the operational mode.

The power in all internally generated spurious signals falling in a bandwidth corresponding to the narrowest spectral channel (approx. 60 kHz) shall be at least 3 dB (TBC) below that of a channel’s nominal noise power.

## Out-of band response

FEND-80930-ZZZ

This section applies only to the operational mode.

For a 2SB mixing scheme, the total output power in the 8-12 GHz IF frequency range shall be at least 10 dB below the total in-band (4-8 GHz) power.

## Saturation

FEND-80940-ZZZ

This section applies only to the operational mode.

For an increase in RF input power of not less than 300 K above TSys, the gain compression shall be less than 1 % for bands 1-6 and less than 3 % for bands 7-10.

## Gain

FEND-80950-ZZZ

This section applies only to the operational mode.

The IF output power spectral density shall be within a range of -40 dBm/GHz to -27 dBm/GHz with an equivalent 300 K RF noise input. This specification is applicable over the full IF band from 4 to 12 GHz.

When the selected band is such that the useful bandwidth is less than 4-12 GHz, the power within 4--12 GHz but outside the useful bandwidth shall be less than -44 dBm (10 dB below the minimum in-band power).

Over the full IF band, the total IF output power shall be less than -18 dBm. and less than -15 dBm over the 10 MHz to 18 GHz frequency range.

The gain variation shall be less than 3 dB (TBC) peak-to-peak over any 2 GHz IF sub-band.  The mean gain-slope shall be less than 6 dB (TBC) over the 4 to 12 GHz IF band.  (This specification permits a maximum 9 dB gain change from one band edge to the other).

Over the useful IF range of the selected band (4 GHz or 8 GHz), and with any LO frequency within the specified range for that band:

a) the gain slope shall be < 5 dB edge-to-edge;

b) the gain ripple shall be < 4.3 dB p-p.

Here "gain slope" is defined to be the edge-to-edge range of the best straight line fit to log[G(f)] where G(f) is the power gain from the front end's input (window) to its IF output as a function of frequency; and "gain ripple" is defined as the peak-to-peak deviation of log[G(f)] from this line.

Over any 2 GHz segment within the useful IF range of the selected band, and with any LO frequency within the specified range for that band:

a) the gain slope shall be < 1.5 dB edge-to-edge;

b) the gain ripple shall be < 1 dB p-p;

c) the difference in gain slope between any two front-end assemblies operating at the same LO frequency shall be < 0.6 dB.

## Gain stability

FEND-80960-ZZZ

This section applies only to the operational mode.

Variations in gain, as defined in section 4.9, as a function of time should comply with the following:

| Application | Power Gain Variation (RMS) |
| --- | --- |
| Single dish observation | 10-4 (TBC) over 1sec |
| Interferometer | 3x10-3 (TBC) over 1 sec |
| Differential polarisation | 5x10-4 (TBC) over 1 sec |

## Signal path phase stability

FEND-80980-ZZZ

The signal path transfer function should maintain a phase stability of better than 1 (TBC) for any 5 minute period for all frequencies within the passband.

# Mechanical and electrical Requirements

## Mass

FEND-93010-ZZZ

The mass of all the Front End components attached to the FESS shall not exceed 750 kg.

Details can be found in the appertaining ICD [AD4].

## Centre of Gravity

FEND-93020-ZZZ

The centre of gravity of all the Front End components attached to the FESS shall be at 400 mm +/- 400 mm below the FESS and within a 250 mm radius of the FESS centre.

## Eigen-frequencies

FEND-93025-ZZZ

The first Eigen-frequencies of any Front End subassembly that is directly attached to any portion of the antenna shall be at least 10 Hz (TBC).

## Volume

FEND-93030-ZZZ

The main Front End components attached to the FESS shall altogether occupy a volume of about 1600 mm x 1000 mm x 1000 mm.

Front End components that will be placed in the volume above the FESS shall not extend beyond a volume of 600 mm x 1380 mm diameter.

The externally mounted compressor (40.03.02.02) shall not extend beyond a volume of 500 mm x 1000 mm x 1050 mm.

The indoor control unit (40.03.02.03) shall not extend beyond a volume of 600 mm x 630 mm x 300 mm.

Details can be found in the appertaining ICD [AD4].

## Orientation

FEND-93040-ZZZ

The Front End shall meet all performance requirements over a range of gravity vectors from 0 to 125 degrees.  This rotation occurs about the antenna elevation-bearing axis.  Details can be found in the appertaining ICD [AD4].

The compressor shall meet all performance requirements at tilt angles up to 10o in any orientation.

## Thermal Load

FEND-93050-ZZZ

The maximum thermal load presented by all the Front End components installed in the receiver cabin during operation shall not exceed 4 kW.  Details can be found in the appertaining ICD [AD4].

## Power requirements

FEND-93060-ZZZ

The Front End components within the receiver cabin shall not consume more than 4 kW of mains-power in the operational mode. The Front End components outside the receiver cabin (compressor) shall not consume more than 10 kW of mains-power. Details can be found in the appertaining ICD [AD4].

The mains power shall be single-phase 230 VAC/50 Hz and three-phase 400 VAC/50 Hz. Details can be found in [AD8].

# operating conditions

## Stabilisation time

FEND-92120-ZZZ

Starting from the non-operational mode, the Front End shall reach the operational mode within 15 minutes meeting all applicable specifications.

100 ms (TBC) after the conclusion of fast-slew, all operational mode specifications shall be met.

## Repeatability

FEND-92125-ZZZ

The performance after a full fast switching cycle with of duration at least 5 minutes shall result in a gain change of less than 2x10-3 (TBC) RMS and a phase change of less than 0.7 (TBC) RMS.

## Simultaneous operation of bands

FEND-92130-ZZZ

This section applies only to the operational and stand-by modes.

Astronomical observations will be performed with one frequency band at any given instant (no dual frequency observations). The water-vapour monitoring radiometer shall operate simultaneously with any of the observing bands.

When band 3 is not used for observation, it shall be always in stand-by mode.  In addition, one additional band can be in stand-by mode.

## Band Selection

This section applies only to the operational and stand-by modes.

## Selection of a (pre-set) observing band

FEND-92140-ZZZ

Selection and operation of a pre-set (i.e. in stand-by mode) observing band shall be possible in less than 1.5 sec. This provides support for fast phase-switching and switching to another (pre-set) observing band.

## Selection of new observing band

FEND-92150-ZZZ

Switching to a new observing band shall be possible in less than 1.5 sec after the new observing band has been switched on twarm-up.  Time twarm-up shall not exceed 15 min. This is to reach thermal equilibrium and maximum stability.

## Narrow-band frequency switching

FEND-92160-ZZZ

Changing between two frequencies within .03% of each other (30 MHz at 100 GHz, 285 MHz at 950 GHz) and in the same band shall require no more than 10 ms.  This provides support for narrow-band frequency switching.

## Frequency changes within a band

FEND-92170-ZZZ

Changing between frequencies more than .03% apart and in the same band shall not require more than 5 s.

## Local Oscillator

This subsection does only apply to the operational mode.

## LO phase stability

FEND-03510-ZZZ

Phase stability (both short and long term – jitter/drift)	63fs and 29 fs  (Note:  The final allocated values are awaiting determination by the Systems Engineering group).

## LO phase resolution

FEND-03310-ZZZ

The first LO shall have a frequency resolution no greater than 0.03 Hz (TBC). The LO shall step and settle to a new frequency in less than 10 ms (TBC) with a phase error of less than 0.6 degrees.

## LO Phase switch step size

FEND-03610-ZZZ

The LO phase shall be switchable from any initial phase in increments of either plus or minus 90 or 180 for periods ranging from 256 us to 16 ms. The final phase value shall settle to within 0.5 (TBC) of the desired value.

## LO Phase switching time

FEND-03620-ZZZ

The final LO phase shall settle to its final value within 1 us after any phase change of 180 degrees or less.

## Monitoring and Control

FEND-43910-ZZZ

All functions in the FE shall be able to be remotely controlled and monitored. Sufficient monitoring shall be provided to indicate the status of the FE and to support preliminary diagnosis for the FE in case of failure. This section applies not for the storage mode. Details can be found in the appertaining ICD [AD6].

## Environmental operating conditions

FEND-92210-ZZZ

## Altitude

The Front End shall meet all its operational performance requirements at an altitude of 5000m.

## Thermal Environment

The Front End shall meet all its operational performance for a ambient temperature of 16o C to 22o C. The temperature variation shall not exceed +/- 1o C peak-to-peak.

In any mode, the Front End shall survive without damage temperatures excursions of -10o C to 50o C (TBC).

## Relative Humidity

The Front End shall meet its performance with a non condensing relative humidity between 20 % (TBC) and 80 %.-

## Vibration

For the transport with the antenna mode, the Front End shall withstand vibrations with a PSD of 1x10-5 g2/Hz from 20-1000 Hz, 6 dB/octave roll-off to 2000 Hz (TBC).

For the transport with the Front End service vehicle, the Front End shall withstand vibrations with a PSD of 1x10-5 g2/Hz from 20-1000 Hz, 6 dB/octave roll-off to 2000 Hz (TBC).

## Acceleration-

For both transport modes and the storage mode, the Front End shall survive without damage accelerations of 4g / -2g for TBD sec.

The maximum slew rate shall be 6 deg/sec and 12 deg/sec2 on both axes for the operational and non-operational mode.

The maximum seismic acceleration shall be 0.3 g on the horizontal axis and 0.3 g on the vertical axis.

## Cleanliness

The Front End shall meet its performance at a cleanliness of TBD.

## Storage and shipping conditions

FEND-92220-ZZZ

This section applies only to the storage mode.

The Front End shall comply with [AD2].

## EMC

FEND-92230-ZZZ

The Front End shall comply with [AD3].

## RFI

FEND-92240-ZZZ

The IF signal isolation between any one operational band and all stand-by bands shall be more than 30 dB.

The RF emission of any cartridge in the 175-191 GHz range shall be at least 10 dB less than the WVR signal level at its maximum IF frequency resolution.

The RF emission of the WVR shall be at least 10 dB less than the IF cartridge signal level at its maximum frequency resolution.

## Grounding / Isolation

FEND-92250-ZZZ

The Front End shall be grounded in compliance with [AD7].

## Availability, reliability and maintainability

FEND-92310-ZZZ

The Front End shall be designed for continuous use. Details can be found in [AD9].

## MTBF

FEND-92320-ZZZ

The MTBF of the Front End shall exceed 5 years (TBC). Details can be found in [AD9].

## MTTR

FEND-92330-ZZZ

At the OSF, the MTTR of the Front End sub-system shall be less than 24 hours (TBC) once all components have reached room temperature.

Exchanging  Front End assemblies at any antenna shall take less than 2 hours (TBC).   Details can be found in [AD9].

## Lifetime

FEND-92340-ZZZ

The lifetime of the Front End shall be class 2 (15 years).

## Maintenance

FEND-92350-ZZZ

There shall be no periodic maintenance of any Front End items, except for some mechanical cryogenic components. Exceptions can be made upon approval from the Front End IPT.

The maintenance interval for the cryo-cooler shall be longer than 10.000 hours.  The maintenance interval for the compressor shall be longer than 20.000 hours.

## Hold times

FEND-92360-ZZZ

After any power interruption not exceeding 30 minutes, all specifications shall be met within 6 hours after the re-application of power.