Procedure for Calculating and Collecting PAI Data for Beam Patterns

ALMA-40.00.00.00-123-A-PRO

Version: A01

Status: Draft

2008-03-10

Prepared By:; Organization; Date. Josh Crabtree; NRAO; 2008-03-10. Approvals:; Organization; Date.

Change Record

| Version | Date | Affected Section(s) | Change Request # | Reason/Initiation/Remarks |
| --- | --- | --- | --- | --- |
| A01 | 2008-03-07 | All | - | First draft |
| A02 | 2008-03-10 | 6.3 |  | Modified to reflect that beam efficiency and phase unwrapping are now performed by the same application. |

Table of Contents

1.	Scope	8

1.1.	Identification	8

1.2.	System overview	8

1.3.	Document overview	8

1.4.	Acronyms and abbreviations	8

2.	Referenced documents	8

2.1.	Applicable documents	8

2.2.	Reference documents	8

3.	Required Data	8

3.1.	All Beam Scans	9

3.2.	Co-Polar Beam Scans	9

3.3.	Cross-Polar Beam Scans	9

4.	Beam Scan Setup	9

5.	Gathering data for all beam scans	11

-Frequency Used For Scan	11

-Elevation (Tilt) Angle	11

-Whether the scan is Co-Polar or Cross-Polar	11

-Whether the scan is for Pol 0 or Pol 1	11

-IF attenuation (dB) used during scan	11

6.	Gathering data for Co-Polar beam scans	11

6.1.	Listings and Images (Script-driven Procedure)	11

6.1.1.	Listings and Images (Script-driven Procedure)	11

6.1.1.1.	Opening and starting the program	11

6.1.1.2.	Select an NSI data file	12

6.1.1.3.	Find the optimal XYZ shift amounts	12

6.1.2.	Find the optimal Z shift amount	13

6.1.3.	Find the optimal X and Y shift amounts	14

6.1.4.	Export Listings and Images	16

6.2.	Listings and Images (Manual Procedure)	17

6.2.1.	Maximum global far field amplitude (dB)	17

6.2.2.	Maximum global near field amplitude (dB)	18

6.2.3.	Shift amounts (m) applied to the far field transform in the X, Y, and Z directions	19

6.2.4.	Hcut (azimuth) and Vcut (elevation) cut images for far field amplitude and phase across subreflector.	23

6.2.5.	Far field amplitude and phase images	24

6.3.	Image of phase on subreflector	27

6.3.1.	Calculate beam center coordinates	27

6.3.2.	Plot Phase Map	29

6.3.3.	If necessary, unwrap phase	30

6.3.3.1.	Simple Unwrap Full Map	31

6.3.3.2.	Unwrap	32

6.3.3.2.1.	Selecting a center point	32

6.3.3.2.2.	Selecting a radius	33

6.3.3.2.2.1.	First Radius Option	34

6.3.3.2.2.2.	Second Radius Option	34

6.3.3.2.2.3.	Third Radius Option	35

6.3.3.2.2.4.	Fourth Radius Option	35

6.3.3.2.2.5.	Fifth Radius Option	35

6.3.3.2.2.6.	Sixth Radius Option	35

6.3.3.2.3.	Number of Iterations	36

6.3.3.2.4.	Unwrap Button	36

6.3.3.2.5.	Specific Approaches to Phase Unwrapping	36

6.3.3.2.5.1.	Approach 1- Simple Unwrap, then Unwrap	36

6.3.3.2.5.2.	Approach 2- Use 360 deg, then 180 deg discontinuity value	37

6.3.3.2.5.3.	Approach 3- Try different center points and/or radii	37

6.3.3.2.5.4.	Approach 4- Find a “Peak Phase Point” through visual inspection	37

6.3.3.2.5.5.	Approach 5- Manually adjust points in a 2-d phase map	41

6.3.3.2.6.	Notes about phase unwrapping	45

6.3.4.	Get Image of Phase Map	45

6.4.	Beam Efficiency Data	45

6.4.1.	Open “Beam Efficiency Calculator.xls”	45

6.4.2.	Select Band Number	45

6.4.3.	Press “Get Center Values” button	45

6.4.4.	Press “Get Beam Efficiencies” button	46

6.5.	Beam center coordinates for near field and hologram	46

6.6.	Plot picture of far field beam center points	47

6.6.1.	Open “Draw Circles and Points.vi” in LabView.	47

6.6.2.	Determine number of points being plotted	47

6.6.3.	Modify block diagram to accommodate the appropriate number of points	47

6.6.4.	Enter beam center coordinates at front panel	49

6.6.5.	Plot the values	50

6.6.6.	Put the appropriate names in the plot legend	51

6.6.7.	Use the same color for all scans at the same frequency	51

6.6.8.	Use circle point style for Pol 0, and square point style for Pol 1	51

6.6.9.	Use hollow point style for Tilt Angle of 0, solid point style for tilt angle of 45.	51

7.	Gathering Data for cross-polar beam scans	51

7.1.	Far field amplitude image	52

8.	Putting the data together	52

8.1.	Open the “Collected Data.xls” spreadsheet	52

8.2.	Create a sheet for each beam scan	52

8.3.	Enter the settings for each beam scan on each sheet	53

8.4.	If Co-Polar, Copy and paste the beam efficiency data into the sheet	53

8.5.	If Co-Polar, Enter the Maximum Global Amplitude and IF Attenuation values for the Far field and near field listings	54

8.6.	If Cross-Polar, Enter the Maximum Global Amplitude and IF Attenuation values for the Far field listing	55

8.7.	If Co-Polar, enter the center values for the near field and hologram listings	55

8.8.	If Co-Polar, enter the XYZ shift amounts	55

8.9.	If Co-Polar, copy and paste the following images into the sheet:	56

8.10.	Enter the appropriate azimuth and elevation values for the nominal beam center at the top of the main page	56

8.11.	Press the “Populate Cells” button on the main page	56

8.12.	Copy and past the picture of the Beam Center plot on the main page	57

8.13.	Important notes regarding the Collected Data sheet	58

A.	Appendix A- “Main Beam Scan Program.bas”	59

B.	Appendix B- “Listings and Images.bas”	64

C.	Appendix C- Nominal Beam Center Coordinates	71

List of Figures

Figure 1- Beam Scan Setup dialog	10

Figure 2- Beam Scan Notes File	10

Figure 3- Opening a Script in NSI2000	12

Figure 4- XYZ Shift Amounts Dialog Window	13

Figure 5- Effect of Z Shift on Phase	14

Figure 6- Phase Cut Across Entire Far field	15

Figure 7- Phase Cut Across Subreflector	15

Figure 8- XYZ Shift Dialog finished	16

Figure 9- Settings for Exported Listings and Images Dialog	16

Figure 10- File Names Displayed in NSI2000 Debug Window	17

Figure 11- Max Global Amplitude in Far Field text file	17

Figure 12- Simple script to create and save Far Field listing	18

Figure 13- Max Global Amplitude in Near Field text file	18

Figure 14- Far Field Button	19

Figure 15- Far field Display Dialog	19

Figure 16- Far Field Transform Setup Dialog	20

Figure 17- Far field Plot Parameters Dialog	20

Figure 18- Hcut and Vcut	21

Figure 19- XYZ Shift Example	22

Figure 20- Far Field Plot Label Centering Off	23

Figure 21- Far Field Transform XYZ Shift Amounts	23

Figure 22- Far Field Span and Center Settings	24

Figure 23- Hcut and Vcut Using XYZ Shift, Centered at Nominal Center Coordinates	24

Figure 24- Far Field display settings for Amplitude and Phase Images	25

Figure 25- Creating a Far Field Amplitude Image	26

Figure 26- Creating a Far Field Phase Image	27

Figure 27- Paste Far Field listing into “Unwrap Phase.xls”	28

Figure 28- Select Band and Press "Get Center Values" Button	28

Figure 29- Plot Circle Encompassing Nominal and Calculated Beam Centers	29

Figure 30- Plot Phase Values Inside Radius Dialog	30

Figure 31- Wrapped Versus Unwrapped Phase Maps	30

Figure 32- Simple Unwrap Full Map Button	31

Figure 33- Discontinuity Value Request Dialog	31

Figure 34- Phase Maps After Using "Simple Unwrap Full Map"	32

Figure 35- Selecting the Appropriate Band	32

Figure 36- Selecting a Center Point	33

Figure 37- Selecting a Radius	34

Figure 38- First Radius Option	34

Figure 39- Second Radius Option	35

Figure 40- third Radius Option	35

Figure 41- Fourth Radius Option	35

Figure 42- Fifth Radius Option	35

Figure 43- Sixth Radius Option	36

Figure 44- Enter the Number of Iterations	36

Figure 45- Unwrap Button	36

Figure 46- Wrapped Phase Map with Sharply Defined Concentric Rings	37

Figure 47- Wrapped Phase. Radius=5, Center= (5,5).	38

Figure 48- Wrapped Phase. Radius=5, Center= (0,0).	39

Figure 49- Wrapped Phase. Radius=5, Center= (-0.5, -0.5).	40

Figure 50- Enter Peak Phase Point values and Press "Get Minimum Radii" button.	41

Figure 51- Phase Map with Small Patch of Wrapped Data Points	42

Figure 52- Data Points Where Wrapping Has Occurred	43

Figure 53- Phase Map After Manually Unwrapping A Small Section	44

Figure 54- Convert Sheet to 1D button	44

Figure 55- Specify Which Sheet Is To Be Copied	45

Figure 56- Select Band Number	45

Figure 57- “Get Center Values” Button	45

Figure 58- "Get Beam Efficiencies" button	46

Figure 61- "Efficiencies" Sheet	46

Figure 62- LabView Diagram Components for a Single Point	47

Figure 63- Block Diagram for Plotting Four Points	48

Figure 64- Block Diagram for Plotting Four Points	49

Figure 65- Front Panel Beam Coordinate Controls	50

Figure 66- Plotted Beam Center Points	50

Figure 67- Plot Name	51

Figure 68- Plot Legend Same Color for Same Frequency	51

Figure 69- Pol 0 with Circles, Pol 1 with Squares	51

Figure 70- Solid points for 45 degree tilt, Hollow points for 0 degree tilt	51

Figure 71- Sheet Names	52

Figure 72- Settings for Beam Scan	53

Figure 73- Paste Efficiency Data Into Spreadsheet	54

Figure 74- Center Points for Near Field and Hologram Listings	55

Figure 75- XYZ Shift Amounts	55

Figure 76- Images Pasted into Sheet	56

Figure 77- Populate Cells Button	57

Figure 78- Snippet of Main Page Data	57

Figure 79- Main Page with Beam Center Plot	58

Figure 80- Blank Sheet for Co-Polar Beam Scan	59

# Scope

# Identification

# System overview

# Document overview

This document is intended to serve as a guide for gathering PAI data pertaining to the testing of a Front End Antenna. Items requiring further consideration or editing are shown in highlighted text. Sections left blank are either not applicable or not yet known. Blank sections should be removed from the final document before approval and release.

# Acronyms and abbreviations

# Referenced documents

# Applicable documents

The following documents are included as part of this document to the extent specified herein. If not explicitly stated differently, the latest issue of the document is valid.

| Reference | Document title | Date | Document ID |
| --- | --- | --- | --- |

# Reference documents

The following documents contain additional information and are referenced in this document.

In the event of a conflict between one of the applicable documents referenced above and the contents of this document, the contents of the applicable document shall be considered as a superseding requirement.

| Reference | Document title | Date | Document ID |
| --- | --- | --- | --- |

# Required Data

The following may be used as a check list for all the necessary data items required for the PAI report.

# All Beam Scans

-Frequency used

-Elevation Angle used

-Whether the scan is Co-Polar or Cross-Polar

-Whether the scan is for Pol 0 or Pol 1

-IF Attenuation (dB) used during scan

-Maximum Global Far Field amplitude (dB)

# Co-Polar Beam Scans

-Maximum Global Near Field amplitude (dB)

-Shift amounts (m) applied to the far field transform in the X, Y, and Z directions

-Beam center coordinates for near field, far field, and hologram

-Plot picture of far field beam center points

-Beam Efficiency data

-Far field amplitude and phase images

-Horizontal (Azimuth) and vertical (Elevation) cut images for far field amplitude and phase across subreflector

-Image of phase on subreflector

# Cross-Polar Beam Scans

-Far field amplitude image

# Beam Scan Setup

Although this document pertains to a procedure performed after a beam scan has already been completed, it is necessary to provide a brief discussion of the setup of a beam scan. Figure 1 shows a dialog presented prior to performing a scan, using the “Main Beam Scan Program.bas” script file. If this file can not be found, the code is provided in Appendix A.

Figure 1- Beam Scan Setup dialog

This dialog window (Figure 1) is presented prior to performing a beam scan. The settings chosen in this dialog window are recorded in a text file which corresponds to the nsi file created during the scan. In this example, the nsi file for the scan is “NRAO FE1 Band3 Jan 2008 Scan 22.nsi”. The notes text file will have the same name as the nsi file, but will have the characters “__NOTES__” appended to the end of the file name. As shown in Figure 2, the notes file contains some of the required data items. The notes file is created in the same directory as the nsi file.

Figure 2- Beam Scan Notes File

# Gathering data for all beam scans

The following five items are found in the scan notes file, as described above in section 4.

# -Frequency Used For Scan

# -Elevation (Tilt) Angle

# -Whether the scan is Co-Polar or Cross-Polar

# -Whether the scan is for Pol 0 or Pol 1

# -IF attenuation (dB) used during scan

# Gathering data for Co-Polar beam scans

In addition to the items listed in section 5, the following items are needed for co-polar scans.

# Listings and Images (Script-driven Procedure)

The steps in this section are all driven by interaction with dialogs in the NSI2000 script environment. The code for the script file (“Listings and Images.bas”) is attached in Appendix B. If the script file can not be found, then the code may be copied and pasted from this document into the NSI2000 scripting utility in order to run. Alternatively, the individual steps may be performed manually, as described below in section 6.2.

# Listings and Images (Script-driven Procedure)

# Opening and starting the program

Click on the Script Button, and then browse to the script file using the file browser in the scripting utility. The green Play button starts the script.

Figure 3- Opening a Script in NSI2000

# Select an NSI data file

A file browser window will prompt for the selection of an NSI data file.

# Find the optimal XYZ shift amounts

A dialog window will assist the user in finding the best XYZ shift amounts. Be sure to initially select the appropriate Band number, and whether Beam 1, Beam 2, or both beams are to be used.

Figure 4- XYZ Shift Amounts Dialog Window

# Find the optimal Z shift amount

The goal is to find the X, Y, and Z shift amounts which produce the flattest phase on the subreflector (at the nominal beam center coordinate). The first step is to increase or decrease the Z shift amount until the phase looks flattest at the nominal center coordinates. After the new Z value is entered, hit the OK button in the dialog, and the program will display either an Azimuth or Elevation phase cut, depending on which was selected in the dialog window. The dialog window will return while the phase cut is visible. The user may continue to adjust the Z shift amount and hit the OK button until the best Z shift amount is determined.

Figure 5- Effect of Z Shift on Phase

# Find the optimal X and Y shift amounts

After the best Z shift amount has been determined, it is necessary to find the best X and Y shift amounts. When finding the X value, make sure to select “View Azimuth Cut”. When finding the Y value, select “View Elevation Cut”. Also, be sure to check the “Zoom In” box. This will make it so that only the phase across the subreflector is displayed (See figures below).

Figure 6- Phase Cut Across Entire Far field

Figure 7- Phase Cut Across Subreflector

The adjustments in the X and Y direction are finer than the Z value shift. Zooming in will make it easier to detect the level of improvement resulting from a new X or Y shift. After the best X and Y values have been determined, it might be useful to tweak the Z shift value a little, and see if the phase improves. After the X, Y and Z shift values have been determined, check the “Check This Box When Finished” box.

Figure 8- XYZ Shift Dialog finished

# Export Listings and Images

After the optimal XYZ shift amounts have been determined, the “Settings for Exported Listings and Images” dialog window will appear. The selected band number, XYZ shift amounts, and beam selection will correspond to the values chosen in the previous dialog, although they may be changed in this dialog if the user wishes to do so. The “Listings” check boxes are used to select which listings will be created. If the “Save Images to Word Doc” box is checked, then a Microsoft Word document will be created containing the following information:

-NSI Scan file name

-X, Y, and Z shift amounts

-Far field listing file header information (Max Global Amplitude, etc.)

-Near field listing file header information (Max Global Amplitude, etc.)

-Far field phase image

-Far field amplitude image

-Hcuts and Vcuts for far field phase on subreflector

Figure 9- Settings for Exported Listings and Images Dialog

The exported files will each contain the original nsi file name, appended by additional text indicating beam selection, XYZ shift amount, and type of file. The files will all be saved to the same directory as the original nsi scan file. Additionally, the names of the newly created files will be displayed in the NSI2000 debug window.

Figure 10- File Names Displayed in NSI2000 Debug Window

# Listings and Images (Manual Procedure)

This section is provided in order to instruct the user on how to manually obtain each of the data items obtained in section 6.1. It is highly recommended that the script-driven procedure of 6.1 be used rather than the manual procedure described in this section. However, it is important to know how to get the data manually, for troubleshooting or other purposes.

# Maximum global far field amplitude (dB)

This item is found in the far field listing created after a scan, using the “Main Beam Scan Program.bas” script file. The far field listing text file will have the same name as the nsi file, except that it is appended with “FF”, as shown in the title bar of the figure below. The maximum global far field amplitude is outlined in red.

Figure 11- Max Global Amplitude in Far Field text file

If the far field listing can not be found, it can be created with the NSI2000 scripting tool, using the code shown in Figure 4. It is first necessary to open the nsi file created during the scan, and then run the script.

Figure 12- Simple script to create and save Far Field listing

# Maximum global near field amplitude (dB)

Figure 13- Max Global Amplitude in Near Field text file

The maximum global near field amplitude is found similarly to the procedure described above for the far field. The near field listing text file will have the same name as the nsi file created during the scan, except that it will end with “NF”. If for some reason the near field listing file can not be found, the near field listing may be created using the code shown in figure 4. In this case, the command “FF_LISTING_TO_FILE” would be replaced with “NF_LISTING_TO_FILE”.

# Shift amounts (m) applied to the far field transform in the X, Y, and Z directions

It is necessary to determine what amounts by which to shift the far field listing in the X, Y, and Z directions. This is accomplished using the following steps.

- Open the nsi file in NSI2000.

- Double click the Far Field button at the top of the screen. This will open the Far Field setup window.

Figure 14- Far Field Button

- Click on each of the three tabs, and make sure the settings are the same as shown in Figures 7, 8, and 9.

Figure 15- Far field Display Dialog

Figure 16- Far Field Transform Setup Dialog

Figure 17- Far field Plot Parameters Dialog

- Click on the “Hcut” and “Vcut” tabs at the top of the screen. This will produce plots of the phase. The Hcut is a plot of the phase at elevation zero degrees, and the Vcut is a plot of the phase at azimuth zero degrees.

Figure 18- Hcut and Vcut

- The above Hcut and Vcut correspond to X, Y, and Z shift amounts of zero meters (shift amounts shown in Figure 16). Experiment with different Z shift amounts until the phase looks the flattest (around the nominal center points). Once the optimal Z shift amount has been determined, start trying different values for the X shift amount. After the optimal X shift amount has been determined, try using different Y shift amounts. It may not be necessary to apply any shift amount in the X or Y direction. The figure below illustrates an example of this process.

Figure 19- XYZ Shift Example

# Hcut (azimuth) and Vcut (elevation) cut images for far field amplitude and phase across subreflector.

- Make sure that plot label centering is turned off.

Figure 20- Far Field Plot Label Centering Off

- Apply optimal XYZ shift amounts, as determined using the procedure from section 6.2

Figure 21- Far Field Transform XYZ Shift Amounts

- Set far field display parameters as shown in the figure below. The Hcut and Vcut images are intended to show the phase across the center of the subreflector. The subreflector has a radius of 3.85 degrees, therefore the span for the images is set to 7.16 degrees (twice the value of the radius). The azimuth and elevation for the center point correspond to the nominal beam center coordinates. In this example for a band 9 scan, the nominal center is 0 (azimuth) and -0.974 (elevation).

Figure 22- Far Field Span and Center Settings

- Press the Hcut and Vcut buttons to create Hcut and Vcut images.

Figure 23- Hcut and Vcut Using XYZ Shift, Centered at Nominal Center Coordinates

# Far field amplitude and phase images

Before creating the far field amplitude and phase images, change the display settings as shown below.

Figure 24- Far Field display settings for Amplitude and Phase Images

- The amplitude image is created by first selecting “Amp” in the far field display dialog, then pressing the “Image” button. Be sure to apply the optimal XYZ shift amounts.

Figure 25- Creating a Far Field Amplitude Image

- The phase image is created similarly, except that “Phase” is selected in the far field display dialog.

Figure 26- Creating a Far Field Phase Image

# Image of phase on subreflector

# Calculate beam center coordinates

Open “Beam Efficiency Calculator.xls”, and paste the far field listing into the four columns at the left hand side of the screen.

Figure 27- Paste Far Field listing into “Unwrap Phase.xls”

Select the appropriate band number, and then press the “Get Center Values” button.

Figure 28- Select Band and Press "Get Center Values" Button

The program will calculate the amplitude center of area for the far field listing. The next step is to plot the phase. The goal is to determine if the phase is unwrapped at two locations:

-Circle with 3.58 degree radius (the radius of the subreflector), centered at the calculate beam center.

-Circle with 3.58 degree radius, centered at the nominal beam center.

After pressing the “Get Center Values” button, the program determines the radius and center coordinates for a circle large enough to encompass the two circles described above.

Figure 29- Plot Circle Encompassing Nominal and Calculated Beam Centers

# Plot Phase Map

After the center values have been obtained, press the “Plot Phase” button.

Figure 30- Plot Phase Values Inside Radius Dialog

When the dialog window appears, type “y” so that only the phase map inside the previously calculated circle will be plotted. If the character “n” is entered, the program will plot the entire phase map. The phase map, in either case, is plotted on Sheet 2.

# If necessary, unwrap phase

“Wrapping” means that there are 360 degree discontinuities in the phase map. After plotting the phase map in the previous step, it is necessary to inspect the phase map for discontinuities. The figure below shows the difference between a wrapped phase map and an unwrapped phase map.

Figure 31- Wrapped Versus Unwrapped Phase Maps

It has been observed that unwrapping will usually be unnecessary if care is taken when determining the best XYZ shift amounts. However, there may be certain beam scans for which unwrapping is unavoidable. The following sections describe the phase unwrapping functions available in the “Beam Efficiency Calculator.xls” application.

# Simple Unwrap Full Map

Figure 32- Simple Unwrap Full Map Button

Figure 33- Discontinuity Value Request Dialog

The “Simple Unwrap Full Map” button causes a simple unwrapping algorithm to be executed across the full phase map. After pressing this button, a dialog will prompt for the size of discontinuity for which to check. The default value is 360, and that will almost always be the only value necessary. There may be occasions when 180 degree discontinuities are present in the phase map after an unwrapping attempt, in which case a value of 180 would be entered. The resulting phase map is plotted on Sheet 4, and the resulting phase listing is copied to the “Phase” column in Sheet 1. This approach will most likely not unwrap the entire phase map completely, but it may have the effect of unwrapping within the circle of interest (Figure 30). In order to verify this, press the “Plot Phase” button to plot the new phase map onto Sheet 2.)

Figure 34- Phase Maps After Using "Simple Unwrap Full Map"

# Unwrap

# Selecting a center point

The first step when selecting a center point is to choose the appropriate band.

Figure 35- Selecting the Appropriate Band

After selecting a band, the nominal center values for that band are copied to the “Nominal Center” field in the “Choose a Center Point” section. The “Calculated Center” field contains the center values obtained after pressing the “Get Center Values” button. The “Peak Phase Point” is initially the same as the calculated center. However, a user may enter different values in this field, which will affect the minimum radii values available (discussed below).

Figure 36- Selecting a Center Point

# Selecting a radius

Figure 37- Selecting a Radius

# First Radius Option

Figure 38- First Radius Option

This option pertains to a circle centered at the Peak Phase Point. The radius is large enough to include a circle (with a radius of 3.58 degrees) centered at the Calculated Center. This option will only be used if the user has entered in a Peak Phase Point that is different from the Calculated Center.

# Second Radius Option

Figure 39- Second Radius Option

This option pertains to a circle centered at the Peak Phase Point. The radius is large enough to include a circle (with a radius of 3.58 degrees) centered at the Nominal Center. This option will only be used if the user has entered in a Peak Phase Point that is different from the Calculated Center.

# Third Radius Option

Figure 40- third Radius Option

This option pertains to a circle centered at the Peak Phase Point. The radius is large enough to include two other circles. The first other circle (with a radius of 3.58 degrees) is centered at the Nominal Center, and the second (also with a radius of 3.58 degrees) is centered at the Calculated Center. This option will only be used if the user has entered in a Peak Phase Point that is different from the Calculated Center.

# Fourth Radius Option

Figure 41- Fourth Radius Option

This option pertains to a circle centered at the Calculated Center Point. The radius is large enough to include a circle (with a radius of 3.58 degrees) centered at the Nominal Center.

# Fifth Radius Option

This option pertains to a circle centered at any of the user-selected center points. The radius is 3.58 degrees, which is the radius of the subreflector.

Figure 42- Fifth Radius Option

# Sixth Radius Option

Figure 43- Sixth Radius Option

This option pertains to a circle centered at the midway point between the Calculated Center and the nominal Center. The radius is large enough to include two other circles. The first other circle (with a radius of 3.58 degrees) is centered at the Nominal Center, and the second (also with a radius of 3.58 degrees) is centered at the Calculated Center.

# Number of Iterations

The last step before pressing the “Unwrap” button is to enter the number of times the algorithm will iterate. This number will almost always be 1, but the user may set it higher if deemed necessary.

Figure 44- Enter the Number of Iterations

# Unwrap Button

Figure 45- Unwrap Button

The “Unwrap” button causes an unwrapping algorithm to be executed only within the user defined circle in the phase map. The user is first asked for the value of discontinuity for which to check, as previously described, after which the unwrapping algorithm is performed. The resulting phase map is plotted on Sheet 4, and the resulting phase listing is copied to the “Phase” column in Sheet 1.

# Specific Approaches to Phase Unwrapping

The process of unwrapping phase can be lengthy and may cause frustration. It is highly recommended that the user spend additional time adjusting the XYZ shift amounts in order to decrease the likelihood that unwrapping will be necessary. However, if unwrapping can not be avoided, the specific approaches may prove helpful. Sometimes a combination of these approaches is necessary.

# Approach 1- Simple Unwrap, then Unwrap

Using this approach, perform the “Simple Unwrap Full Map” step as described above. Afterward, copy the “Unwrapped Phase” column into the “Phase” column. Then press the “Unwrap” button. This approach is useful if the original phase map was wrapped in numerous places within the circle of interest, and the “Simple Unwrap” was able to reduce the amount of wrapping.

# Approach 2- Use 360 deg, then 180 deg discontinuity value

This approach may be useful if an initial unwrapping attempt appeared to be almost successful, but has caused 180 degree discontinuities to remain in the map.

# Approach 3- Try different center points and/or radii

# Approach 4- Find a “Peak Phase Point” through visual inspection

Sometimes a phase map will contain wrapping in the form of sharply defined concentric rings. In this approach, the center point of the concentric rings is determined, and used as the center for unwrapping.

Figure 46- Wrapped Phase Map with Sharply Defined Concentric Rings

- The first step in this approach is to set the Plot Center to (5, 5), and set the radius to 5 degrees. Then hit the “Plot Phase” button, and type “y” at the prompt asking if only the phase inside the radius should be plotted.

Figure 47- Wrapped Phase. Radius=5, Center= (5,5).

The result is plotted on Sheet 2. Orient the plot so that the phase circle is located at the upper right hand corner.

- The next step is to select a Plot Center location of (0, 0), and then plot the phase again.

Figure 48- Wrapped Phase. Radius=5, Center= (0,0).

- At this point start adjusting the center coordinates until the phase map appears to be centered at the center of the concentric rings. In this example, the center was determined to be located at (-0.5, -0.5).

Figure 49- Wrapped Phase. Radius=5, Center= (-0.5, -0.5).

- After the peak phase point has been determined, enter this value into the “Plot Center” section as well as the “Peak Phase Point” section.

Figure 50- Enter Peak Phase Point values and Press "Get Minimum Radii" button.

- The next step is to select the desired radius, and proceed with unwrapping according to any of the procedures previously described.

# Approach 5- Manually adjust points in a 2-d phase map

This approach should generally be avoided. However, if unwrapping attempts continually result in a phase map that is completely unwrapped except for a very small number of data points, this approach may be the only viable option. In this example, a phase map has a small set of data points which protrude above the surrounding data.

Figure 51- Phase Map with Small Patch of Wrapped Data Points

Close inspection of the data reveals that the values appear to be 360 degrees higher than the immediately surrounding points.

Figure 52- Data Points Where Wrapping Has Occurred

The solution is to simply subtract 360 from each of the data points in the wrapped region (shaded in yellow in Figure 51). It should be noted that the discontinuity may be larger or smaller than 360, depending on how many attempts at unwrapping have been performed. It is best to first determine the size of the discontinuity at one of the points. After subtracting (or adding, if required) the value of discontinuity at each point, it is necessary to inspect the plot image again to verify that the unwrapping has been successful.

Figure 53- Phase Map After Manually Unwrapping A Small Section

If the result looks satisfactory, it is then necessary to copy the contents of the current sheet into the “Phase” column on Sheet 1. This is achieved by first pressing the “Convert Sheet to 1D” button, and then specifying which sheet is to be copied.

Figure 54- Convert Sheet to 1D button

Figure 55- Specify Which Sheet Is To Be Copied

# Notes about phase unwrapping

      - If an attempt at phase unwrapping is successful, it is recommended that the file be saved using a name something like “Phase Unwrap Band 9 Scan 58.xls”.

# Get Image of Phase Map

The image of the phase will be displayed on either Sheet 2 or Sheet 4, depending on which steps were used for plotting and/or unwrapping.

# Beam Efficiency Data

# Open “Beam Efficiency Calculator.xls”

This file should already be open after the previous step. The four columns at the left hand side (Azimuth, Elevation, Amplitude, Phase) should contain the far field listing after the previous step.

# Select Band Number

Figure 56- Select Band Number

# Press “Get Center Values” button

Figure 57- “Get Center Values” Button

# Press “Get Beam Efficiencies” button

Figure 58- "Get Beam Efficiencies" button

This cause the program to calculate beam efficiency values. The resulting data will be displayed on the “Efficiencies” sheet.

Figure 59- "Efficiencies" Sheet

# Beam center coordinates for near field and hologram

In order to get the beam center coordinates for the near field, paste the Nearfield listing into the “Beam Efficiency Calculator.xls” sheet. Then press the “Get Center Values” button. The calculated center will be displayed in three locations- the “Calculated Center” fields, the “Plot Center” fields, and the “Peak Phase Point” fields. Repeat this process with the hologram listing to get the beam center for the hologram.

# Plot picture of far field beam center points

After the efficiency data has been calculated for all of the co-polar beam scans, it is necessary to plot the locations of the calculated far field beam centers with respect to the nominal beam center. The center values may either be obtained from the “Phase Unwrap.xls” sheet, or the beam efficiency data described in section 6.4.

# Open “Draw Circles and Points.vi” in LabView.

There will be a point plotted for each co-polar beam scan.

# Determine number of points being plotted

There will be a point plotted for each co-polar beam scan.

# Modify block diagram to accommodate the appropriate number of points

Go to the block diagram. For each point being plotted, there will be this diagram:

Figure 60- LabView Diagram Components for a Single Point

So, for example, if four points are to be plotted, the diagram will look like this:

Figure 61- Block Diagram for Plotting Four Points

If only two points are to be plotted, the block diagram will look like this:

Figure 62- Block Diagram for Plotting Four Points

# Enter beam center coordinates at front panel

The front panel has a section for beam center coordinates. The two controls at the top, “AZ Nominal” and “EL Nominal”, will have azimuth and elevation values for the nominal center. The rest of these controls will contain the azimuth and elevation values for each of the calculated beam centers.

Figure 63- Front Panel Beam Coordinate Controls

# Plot the values

Hit the start button in LabView, and the points will be plotted in the graph.

Figure 64- Plotted Beam Center Points

# Put the appropriate names in the plot legend

Figure 65- Plot Name

The name should refer to the frequency, pol number, and tilt angle used in the scan.

# Use the same color for all scans at the same frequency

Figure 66- Plot Legend Same Color for Same Frequency

# Use circle point style for Pol 0, and square point style for Pol 1

Figure 67- Pol 0 with Circles, Pol 1 with Squares

# Use hollow point style for Tilt Angle of 0, solid point style for tilt angle of 45.

Figure 68- Solid points for 45 degree tilt, Hollow points for 0 degree tilt

There may be more than two tilt angles among the scans being used. If that is the case, there are other ways to differentiate the appearance of the points.

# Gathering Data for cross-polar beam scans

In addition to the items listed in section 5, the following item is needed for cross-polar scans.

# Far field amplitude image

This image is obtained using the procedure described above in section 6.5.

# Putting the data together

After the data has all been collected, it is time to put it all together.

# Open the “Collected Data.xls” spreadsheet

# Create a sheet for each beam scan

The name of the sheet should contain frequency, Pol number, Co-Pol or X-Pol, and Elevation angle.

Figure 69- Sheet Names

The sheets should be arranged by frequency, sub arranged by tilt angle, further sub arranged by Pol number, and Co-Pol/Cross. For example, if 12 scans were performed, where the first 8 were at 200 GHz, and the last four were at 500 GHz, the sheet order would look something like this:

- 200 GHz, Pol 0, Co-Pol, Elevation 0

- 200 GHz, Pol 0, Cross-Pol, Elevation 0

- 200 GHz, Pol 1, Co-Pol, Elevation 0

- 200 GHz, Pol 1, Cross-Pol, Elevation 0

- 200 GHz, Pol 0, Co-Pol, Elevation 45

- 200 GHz, Pol 0, Cross-Pol, Elevation 45

- 200 GHz, Pol 1, Co-Pol, Elevation 45

- 200 GHz, Pol 1, Cross-Pol, Elevation 45

- 500 GHz, Pol 0, Co-Pol, Elevation 0

- 500 GHz, Pol 0, Cross-Pol, Elevation 0

- 500 GHz, Pol 1, Co-Pol, Elevation 0

- 500 GHz, Pol 1, Cross-Pol, Elevation 0

The above order is recommended, but it is not necessary if there aren’t a complete set of scans for a certain frequency. There also may be more than one scan at exactly the same settings. In this case it is good to put the extra scans at the end of the set with the same frequency. For example, if there were a complete set of four scans done at a tilt angle of 45 degrees, and then two more scans were done at the same setting as one the first scan, the order would look something like this:

- 200 GHz, Pol 0, Co-Pol, Elevation 0

- 200 GHz, Pol 0, Cross-Pol, Elevation 0

- 200 GHz, Pol 1, Co-Pol, Elevation 0

- 200 GHz, Pol 1, Cross-Pol, Elevation 0

- 200 GHz, Pol 0, Co-Pol, Elevation 0 (second scan at this setting)

- 200 GHz, Pol 0, Co-Pol, Elevation 0 (third scan at this setting)

# Enter the settings for each beam scan on each sheet

The settings for each beam scan will be at the top left corner of each sheet. Fill in the fields with the appropriate information for the scan. The information for this section was obtained as described in Section 4 (see also Figure 2).

Figure 70- Settings for Beam Scan

# If Co-Polar, Copy and paste the beam efficiency data into the sheet

This data was obtained as described in Section 6.4 (see also Figure 61).

Figure 71- Paste Efficiency Data Into Spreadsheet

# If Co-Polar, Enter the Maximum Global Amplitude and IF Attenuation values for the Far field and near field listings

The IF attenuation value was obtained as described in Section 5 (see also Figure 2). The Max Global map values are found in the Microsoft Word document created using the script-driven procedure, as described in Section 6.1.4. If the script-driven procedure was not used, then the Max Global Amp values were obtained as described in Sections 6.2.1 and 6.2.2 (see also Figures 11 and 13).

# If Cross-Polar, Enter the Maximum Global Amplitude and IF Attenuation values for the Far field listing

This is the same as was done for the co-polar scan, but the near field data is not required.

# If Co-Polar, enter the center values for the near field and hologram listings

This data was obtained as described in Section 6.5.

Figure 72- Center Points for Near Field and Hologram Listings

# If Co-Polar, enter the XYZ shift amounts

These values were obtained as described in Section 6.1.3 (script-driven procedure) and Section 6.2.3 (manual procedure).

Figure 73- XYZ Shift Amounts

# If Co-Polar, copy and paste the following images into the sheet:

- Far field amplitude

- Far field phase

- Far field Hcuts (amplitude and phase)

- Far field Vcuts (amplitude and phase)

- Far Field Phase on subreflector

This information was obtained according to the procedures described in Sections 6.1.4, 6.2.4, 6.2.5 and 6.3.

Figure 74- Images Pasted into Sheet

# Enter the appropriate azimuth and elevation values for the nominal beam center at the top of the main page

# Press the “Populate Cells” button on the main page

Figure 75- Populate Cells Button

Once all the sheets have been filled in and the nominal beam center values have been entered, press the “Populate Cells” button on the main page. This will collect and calculate all of the data needed for the main page.

Figure 76- Snippet of Main Page Data

# Copy and past the picture of the Beam Center plot on the main page

Figure 77- Main Page with Beam Center Plot

# Important notes regarding the Collected Data sheet

Make sure to paste everything into the correct cells in the template sheets. The “Collected Data.xls” spreadsheet has blank template sheets for co-polar and cross-polar scans. It is important to follow these templates. The “Populate Cells” button triggers a macro which reads data from specific cells in each sheet. If any of the data is in the wrong cell, then it won’t be read.

Figure 78- Blank Sheet for Co-Polar Beam Scan

# Appendix A- “Main Beam Scan Program.bas”

The following code is used to perform beam scans and create far field, near field, and hologram listings after each scan. If the script file can not be found, this code may be copied and pasted into the NSI2000 scripting utility and run.

'********************

'* Main Scan Program

'********************

Sub Main

On Error Goto labelEnd

Debug.Clear

Dim Zposition As String

Dim BeginningScanNumber As Integer

Dim NumberOfScans As Integer

Dim filenumber As String

Dim filename1 As String

Dim filename2 As String

Dim count As Integer

Dim Zpos As Double

Dim AxisName As String

Dim lZ As Long

Dim k As Integer

Dim FileNamePrefix As String

Dim PrefixAndNumber As Variant

Dim tmpName As String

Dim NoteFileName As String

'********************************************************

' Get information to make file names.

'********************************************************

tmpName = DATA_FILENAME

PrefixAndNumber = GetPrefixAndNumber(tmpName)

FileNamePrefix = PrefixAndNumber(0)

BeginningScanNumber = CInt(PrefixAndNumber(1))

'********************************************************

' Get current value for Z distance for Beam 1.

' Move Z axis to that position.

'********************************************************

'find beam table column index for Z-Axis

For k = 1 To BeamTable.AxesCount

If BeamTable.AxisName(CLng(k)) = "Z axis" Then

lZ = k

End If

Next k

'get Z-Axis position for first beam

Zposition = CStr(BeamTable.Item(1).PositionArray(lZ))

Debug.Print "Z position " & Zposition

MOVE_TO_SCAN_CENTER

CONTROLLER_AUX_AXIS_SET("Z axis")

CONTROLLER_AUX_AXIS_MOVE Zposition

BeginningScanNumber = CInt(PrefixAndNumber(1))

'Debug.Print "Beginning Scan number = " & BeginningScanNumber

'********************

' Perform Scans

'********************

NumberOfScans = InputBox("", "Enter the number of Scans")

If (CStr(NumberOfScans) = "" Or NumberOfScans = 0) Then

GoTo labelEnd

End If

NoteFileName = CreateTextFile(NumberOfScans, FileNamePrefix, BeginningScanNumber)

Debug.Print "Performing " & NumberOfScans & " scans"

BeginningScanNumber = BeginningScanNumber - 1

'Debug.Print "Beginning Scan number = " & BeginningScanNumber

count = 0

For I = (BeginningScanNumber+1) To (BeginningScanNumber + NumberOfScans)

filename1 = FileNamePrefix + CStr(I) + ".nsi"

CONTROLLER_AUX_AXIS_SET("Z axis")

CONTROLLER_AUX_AXIS_MOVE Zposition

MOVE_TO_SCAN_CENTER

CONTROLLER_AUX_AXIS_SET("Z axis")

CONTROLLER_AUX_AXIS_MOVE Zposition

'Important: Leave these two lines exactly as they are

MEAS_CREATE_NEW_SCAN

MEAS_ACQUIRE filename1, True

CONTROLLER_AUX_AXIS_SET "Z axis"

CONTROLLER_AUX_AXIS_MOVE Zposition

Debug.Print filename1

count = count + 1

Next I

Debug.Print "Scans completed."

Debug.Print vbCrLf

'*******************************

' Create Far-Field Listings

'*******************************

For J = (BeginningScanNumber+1) To (BeginningScanNumber + NumberOfScans)

filename1 = FileNamePrefix + CStr(J) + ".nsi"

filename2 = FileNamePrefix + CStr(J) + " FF.txt"

DATA_FILENAME = filename1

If FF_PLOTCENTERING = 1 Then

FF_PLOTCENTERING_OFF

End If

SELECT_BEAM 1

SELECT_BEAM 2, 1

FF_LISTING_TO_FILE filename2

Debug.Print filename2

Next J

Debug.Print "FF listings completed"

'*******************************

' Create Near-Field Listings

'*******************************

For J = (BeginningScanNumber+1) To (BeginningScanNumber + NumberOfScans)

filename1 = FileNamePrefix + CStr(J) + ".nsi"

filename2 = FileNamePrefix + CStr(J) + " NF.txt"

DATA_FILENAME = filename1

If NF_PLOTCENTERING = 1 Then

NF_PLOTCENTERING_OFF

End If

SELECT_BEAM 1

SELECT_BEAM 2, 1

NF_LISTING_TO_FILE filename2

Debug.Print filename2

Next J

Debug.Print "NF listings completed"

'*******************************

' Create Hologram Listings

'*******************************

For J = (BeginningScanNumber+1) To (BeginningScanNumber + NumberOfScans)

filename1 = FileNamePrefix + CStr(J) + ".nsi"

filename2 = FileNamePrefix + CStr(J) + " HOLO.txt"

DATA_FILENAME = filename1

If HOLO_PLOTCENTERING = 1 Then

HOLO_PLOTCENTERING_OFF

End If

SELECT_BEAM 1

SELECT_BEAM 2, 1

HOLO_LISTING_TO_FILE filename2

Debug.Print filename2

Next J

Debug.Print "HOLOGRAM listings completed"

Debug.Print "Notes saved to file: " & NoteFileName

'Open the last nsi file created

DATA_FILENAME = filename1

labelEnd:

End Sub

Function GetPrefixAndNumber(tmp As String) As Variant

Dim lengthSuffix As Long

Dim tmpchar As String

Dim count As Integer

Dim done As Boolean

Dim AscVal As Integer

Dim ResultArray(2) As String

Dim prefix As String

Dim ScanNumber As String

count = 0

done = False

lengthSuffix = 4

Do While done = False

tmpchar = Mid(tmp, Len(tmp) - 4 - count, 1)

AscVal = Asc(tmpchar)

lengthSuffix = lengthSuffix + 1

If (AscVal <48) Or (AscVal > 57) Then

done = True

lengthSuffix = lengthSuffix - 1

End If

count = count + 1

Loop

prefix = Left(tmp, Len(tmp)-lengthSuffix)

ScanNumber = Left(Right(tmp, lengthSuffix), lengthSuffix - 4)

ResultArray(0) = prefix

ResultArray(1) = ScanNumber

GetPrefixAndNumber = ResultArray

End Function

Function CreateTextFile(NumberOfScans As Integer, FileNamePrefix As String, FirstScanNumber As Integer) As String

Dim notes As String

Dim filename As String

Dim tmpstr As String

Dim tmpchar As String

Dim done As Boolean

Dim filenum As Integer

Dim count As Long

Dim lengthEND As Long

Dim CoCross As String

Dim PolNum As String

Dim FirstScanNum As Integer

FirstScanNumber = FirstScanNumber + 1

Begin Dialog UserDialog 550,287,"Notes" ' %GRID:10,7,1,1

TextBox 150,42,100,21,.AttenValue

TextBox 150,70,100,21,.TiltAngleValue

OptionGroup .Group1

OptionButton 10,98,20,14,"CoPolButton",.OptionButton1

OptionButton 10,119,20,14,"CrossPolButton",.CrossPolButton

Text 40,98,80,14,"Co-Polar",.Text1

Text 40,119,100,14,"Cross-Polar",.Text2

TextBox 150,14,100,21,.FrequencyValue

Text 10,21,130,21,"RF Frequency (GHz)",.Text3

Text 10,42,130,21,"IF Attenuation (dB)",.Text4

Text 10,70,130,21,"Tilt Angle (deg)",.Text5

TextBox 370,14,100,21,.XcenterValue

Text 300,21,70,14,"X Center",.Text6

TextBox 370,42,100,21,.YcenterValue

Text 300,49,60,14,"Y Center",.Text7

OptionGroup .Group2

OptionButton 160,98,20,14,"OptionButton2",.Pol0Button

OptionButton 160,119,20,14,"Pol1Button",.OptionButton3

Text 190,98,60,14,"Pol0",.Text8

Text 190,119,50,14,"Pol1",.Text9

TextBox 10,161,530,84,.AdditionalNotesValue

Text 10,147,240,14,"Additional Notes",.Text10

OKButton 20,259,120,21

CancelButton 160,259,120,21

End Dialog

Dim dlg As UserDialog

dlg.XcenterValue = CStr(Round(MEAS_HCENTER, 5))

dlg.YcenterValue = CStr(Round(MEAS_VCENTER, 5))

dlg.FrequencyValue = MEAS_FREQUENCY

dlg.AttenValue = "0"

dlg.TiltAngleValue = "0"

dlg.AdditionalNotesValue = "None"

Dialog dlg

If dlg.group1 = 0 Then

CoCross = "Co-Polar"

End If

If dlg.group1 = 1 Then

CoCross = "Cross-Polar"

End If

If dlg.group2 = 0 Then

PolNum = "Pol 0"

End If

If dlg.group2 = 1 Then

PolNum = "Pol 1"

End If

SELECT_BEAM 1

notes = "RF Frequency: " & dlg.FrequencyValue & vbCrLf

notes = notes & "IF Attenuation (dB)= "  & dlg.AttenValue & vbCrLf

notes = notes & "Tilt Angle (deg): " & dlg.TiltAngleValue & vbCrLf

notes = notes & CoCross & vbCrLf

notes = notes & PolNum & vbCrLf

notes = notes & "X center: " & dlg.XcenterValue & ", Y center: " & dlg.YcenterValue

notes = notes & vbCrLf & vbCrLf & dlg.AdditionalNotesValue

'MsgBox (notes, vbOkOnly, "Notes")

tmpstr = DATA_FILENAME

done = False

count = 1

Do While done = False

tmpchar = Mid(tmpstr, Len(tmpstr)-count, 1)

count = count + 1

If tmpchar = "\" Then

done = True

count = count - 1

End If

lengthEND = count

Loop

'filename = Left(tmpstr, Len(tmpstr)-lengthEND) & "NOTES__" & Left(Right(tmpstr, lengthEND), lengthEND - 4) & ".txt"

filename = FileNamePrefix & CStr(FirstScanNumber) & "__NOTES__.txt"

'MsgBox "Filename is " & filename

CreateTextFile = filename

Debug.Print "File name = " & filename

filenum = FreeFile

Open filename For Output As #filenum

Print #filenum, "Date/Time: " & Now & vbCrLf

Print #filenum, "Last scan was: " & tmpstr & vbCrLf

Print #filenum, "Number of scans in this series: " & NumberOfScans & vbCrLf

Print #filenum, "Notes:" & vbCrLf & notes

Close #filenum

End Function

# Appendix B- “Listings and Images.bas”

The following code allows a user to interact with a dialog to find the best XYZ shfit amounts. The program will also create far field, near field, and hologram listings with user-selected shifts in the X, Y, and Z directions, and export amplitude and phase images to a Microsoft Word document. If the script file can not be found, this code may be copied and pasted into the NSI2000 scripting utility and run.

'*************************************************************

'User interacts with dialog to select best X, Y, and Z shift

'amounts. Then, the NF, FF, and HOLO listings are exported as

'text files, and the images are saved to a Microsoft Word doc.

'*************************************************************

Sub Main

On Error GoTo labelEnd

Dim filename As String

Dim filenameFF As String

Dim filenameNF As String

Dim filenameHOLO As String

Debug.Clear

Dim tmp As String

Dim NominalCenterAZ As Double

Dim NominalCenterEL As Double

Dim Band As Integer

Dim Xshift As Double

Dim Yshift As Double

Dim Zshift As Double

Dim Beams As String

Dim VarTamp As Variant

tmp = GetFilePath("", "nsi", "C:\NSI2000\Data" , "Please select an NSI data file")

If tmp = "" Then

GoTo labelEnd

End If

DATA_FILENAME = tmp

filename = DATA_FILENAME

MEAS_PROBE_MODEL_NONE

varTmp = GetOptimalXYZ()

Begin Dialog UserDialog 860,182,"Settings for Exported Listings and Images" ' %GRID:10,7,1,1

GroupBox 20,63,170,112,"Please Select Band",.GroupBox1

OptionGroup .Group1

OptionButton 30,84,90,14,"Band 3",.ButtonBand3

OptionButton 30,105,90,14,"Band 6",.ButtonBand6

OptionButton 30,126,80,14,"Band 7",.ButtonBand7

OptionButton 30,147,80,14,"Band 9",.ButtonBand9

GroupBox 10,7,840,42,"Data File",.GroupBox2

OKButton 700,70,130,28

CancelButton 700,105,130,28

GroupBox 210,63,130,112,"Listings",.GroupBox3

CheckBox 220,84,100,14,"Far Field",.CheckBoxFF

CheckBox 220,105,90,14,"Near Field",.CheckBoxNF

CheckBox 220,126,110,14,"Hologram",.CheckBoxHOLO

GroupBox 350,63,170,112,"XYZ Shift Amounts (m)",.GroupBox4

Text 359,88,20,14,"X:",.Text1

TextBox 380,84,100,21,.TextBoxXshift

Text 360,115,20,14,"Y:",.Text2

TextBox 380,112,100,21,.TextBoxYshift

Text 360,143,20,14,"Z:",.Text3

TextBox 380,140,100,21,.TextBoxZshift

GroupBox 530,63,150,84,"Beam Select",.GroupBox5

OptionGroup .Group2

OptionButton 540,84,110,14,"Beam 1",.OptionButton1

OptionButton 540,105,80,14,"Beam 2",.OptionButton2

OptionButton 540,126,100,14,"Both Beams",.OptionButton3

CheckBox 530,154,200,14,"Save Images to Word Doc",.CheckBoxSaveImages

Text 20,21,820,21,filename,.TextFileName

End Dialog

Dim dlg As UserDialog

dlg.CheckBoxFF = 1

dlg.CheckBoxNF = 1

dlg.CheckBoxHOLO = 1

dlg.CheckBoxSaveImages = 1

dlg.group1 = vartmp(0)

dlg.Group2 = vartmp(1)

dlg.TextBoxXshift = FF_XSHIFT

dlg.TextBoxYshift = FF_YSHIFT

dlg.TextBoxZshift = FF_ZSHIFT

Dialog dlg

Select Case dlg.group1

Case 0

Band = 3

NominalCenterAZ = 0.323

NominalCenterEL = 1.811

Case 1

Band = 6

NominalCenterAZ = 1.6867

NominalCenterEL = -1.6867

Case 2

Band = 7

NominalCenterAZ = 0.974

NominalCenterEL = 0

Case 3

Band = 9

NominalCenterAZ = 0

NominalCenterEL = -0.974

End Select

Select Case dlg.group2

Case 0

SELECT_BEAM 1

Beams = "Beam1"

Case 1

SELECT_BEAM 2

Beams = "Beam2"

Case 2

SELECT_BEAM 1

SELECT_BEAM 2,1

Beams = "Both Beams"

End Select

Xshift = CDbl(dlg.TextBoxXshift)

Yshift = CDbl(dlg.TextBoxYshift)

Zshift = CDbl(dlg.TextBoxZshift)

FF_XSHIFT = Xshift

FF_YSHIFT = Yshift

FF_ZSHIFT = Zshift

FF_HCENTER = 0

FF_VCENTER = 0

FF_HSPAN = 20

FF_VSPAN = 20

If dlg.CheckBoxHolo = 1 Then

'Create Hologram file

filenameHOLO = Fname(filename, "HOLO", ".txt", Beams)

Debug.Print "Hologram Listing : " & filenameHOLO

HOLO_HSPAN = 0.1

HOLO_VSPAN = 0.1

HOLO_HCENTER = 0

HOLO_VCENTER = 0

HOLO_HPTS = 200

HOLO_VPTS = 200

HOLO_PLOTCENTERING_OFF

HOLO_HVFFTMIN 128, 128

HOLO_NumberDecimalDigitsInListing = 4

HOLO_LISTING_TO_FILE filenameHOLO, False

End If

If dlg.CheckBoxNF = 1 Then

'Create NF file

filenameNF = Fname(filename, "NF", ".txt", Beams)

Debug.Print "Nearfield Listing: " & filenameNF

NF_PLOTCENTERING_OFF

NF_NumberDecimalDigitsInListing = 4

NF_LISTING_TO_FILE filenameNF, False

End If

If dlg.CheckBoxFF = 1 Then

'Create FF file

filenameFF = Fname(filename, "FF", ".txt", Beams)

Debug.Print "Farfield Listing : " & filenameFF

FF_HSPAN = "20 deg"

FF_VSPAN = "20 deg"

FF_HCENTER = 0

FF_VCENTER = 0

FF_HPTS = 201

FF_VPTS = 201

FF_PLOTCENTERING_OFF

FF_NumberDecimalDigitsInListing = 4

FF_LISTING_TO_FILE filenameFF, False

End If

If dlg.CheckBoxSaveImages = 1 Then

Dim Wapp As Object

Dim Wdoc As Object

Set Wapp = CreateObject("word.application")

Wapp.Visible = False

Set Wdoc = Wapp.Documents.Add()

Wapp.Selection.TypeText filename & vbCrLf & vbCrLf

Wapp.Selection.TypeText "X shift (m): " & Xshift & vbCrLf

Wapp.Selection.TypeText "Y shift (m): " & Yshift & vbCrLf

Wapp.Selection.TypeText "Z shift (m): " & Zshift & vbCrLf & vbCrLf & vbCrLf

FF_PC_NONE

MEAS_PROBE_MODEL_NONE

FF_PLOTCENTERING_OFF

FF_AMPLITUDE

FF_IMAGE

Wapp.Selection.TypeText PLOT_BLOCK_FF_MAX

CLEAR_ALL_PLOTS

MEAS_PROBE_MODEL_NONE

NF_PLOTCENTERING_OFF

NF_AMPLITUDE

NF_IMAGE

Wapp.Selection.TypeText PLOT_BLOCK_NF_MAX

CLEAR_ALL_PLOTS

MEAS_PROBE_MODEL_NONE

FF_PLOTCENTERING_OFF

FF_AMPLITUDE

FF_IMAGE

COPY_LAST_PLOT

Wapp.Selection.Paste

CLEAR_ALL_PLOTS

MEAS_PROBE_MODEL_NONE

FF_PHASE

FF_IMAGE

COPY_LAST_PLOT

Wapp.Selection.Paste

CLEAR_ALL_PLOTS

'Use Nominal center to view Hcuts and Vcuts

FF_HCENTER = NominalCenterAZ

FF_VCENTER = NominalCenterEL

FF_HSPAN = 8

FF_VSPAN = 8

MEAS_PROBE_MODEL_NONE

FF_AMPLITUDE

FF_HCUT

COPY_LAST_PLOT

Wapp.Selection.Paste

CLEAR_ALL_PLOTS

FF_VCUT

COPY_LAST_PLOT

Wapp.Selection.Paste

CLEAR_ALL_PLOTS

MEAS_PROBE_MODEL_NONE

FF_PHASE

FF_HCUT

COPY_LAST_PLOT

Wapp.Selection.Paste

CLEAR_ALL_PLOTS

FF_VCUT

COPY_LAST_PLOT

Wapp.Selection.Paste

CLEAR_ALL_PLOTS

Wdoc.SaveAs Fname(filename, "IMAGES", ".doc", Beams)

Wdoc.Close

Set Wapp = Nothing

Set Wdoc = Nothing

Debug.Print "Images saved to  : " & Fname(filename, "IMAGES", ".doc", Beams)

End If

CLEAR_ALL_PLOTS

FF_HSPAN = 20

FF_VSPAN = 20

FF_HCENTER = 0

FF_VCENTER = 0

labelEnd:

End Sub

Function Fname(filename As String, filetype As String, Extension As String, Beams As String) As String

Dim length As Long

length = Len(filename) - 4

Dim ShiftString As String

ShiftString = " Xshift_" & FF_XSHIFT & " Yshift_" & FF_YSHIFT & " Zshift_" & FF_ZSHIFT

Fname = Left(filename, length) & " " & filetype & " " & Beams & ShiftString & Extension

End Function

Function GetOptimalXYZ() As Variant

Dim count As Integer

count = 1

Dim done As Boolean

Dim ZoomIn As Boolean

Dim BandSelect As Integer

Dim NominalCenterAZ As Double

Dim NominalCenterEL As Double

Dim AzCut As Boolean

done = False

On Error GoTo ExitLoop

CLEAR_ALL_PLOTS

ZoomIn = False

Dim ResultsArray(4) As Double

FF_PHASE

FF_PC_NONE

'Initalize to band 3 settings, let user choose otherwise in dialog

BandSelect = 0

NominalCenterAZ = 0.323

NominalCenterEL = 1.811

FF_XSHIFT = 0

FF_YSHIFT = 0

FF_ZSHIFT = 0

AzCut = False

SELECT_BEAM 1

'SELECT_BEAM 2, 1

Begin Dialog UserDialog 454,274,"Select Best XYZ Shift Amounts" ' %GRID:10,7,1,0

GroupBox 15,0,340,60,"",.GroupBox1

Text 20,28,100,14,"Nominal Center",.Text1

Text 143,10,90,14,"Azimuth",.Text2

Text 254,11,90,14,"Elevation",.Text3

Text 23,78,22,15,"X",.Text4

TextBox 50,77,80,21,.TextBoxX

Text 22,101,18,15,"Y",.Text5

Text 24,123,20,18,"Z",.Text6

TextBox 50,100,79,17,.TextBoxY

TextBox 48,122,79,18,.TextBoxZ

GroupBox 20,144,116,97,"",.GroupBox2

OKButton 299,156,92,23

OptionGroup .Group2

OptionButton 30,154,80,21,"Band 3",.OptionButtonBand3

OptionButton 30,175,80,21,"Band 6",.OptionButtonBand6

OptionButton 30,196,80,21,"Band 7",.OptionButtonBand7

OptionButton 30,217,100,21,"Band 9",.OptionButtonBand9

CancelButton 299,183,92,25

GroupBox 162,70,268,73,"",.GroupBox3

OptionGroup .Group1

OptionButton 170,77,180,21,"View Azimuth Cut",.OptionButtonViewAZcut

OptionButton 170,98,180,21,"View Elevation Cut",.OptionButtonViewELcut

CheckBox 171,120,250,18,"Zoom In (View only on subreflector)",.CheckBoxZoomIn

TextBox 128,28,90,18,.TextBoxNomCtrAZ

TextBox 239,27,95,20,.TextBoxNomCtrEL

CheckBox 151,225,231,23,"Check This Box When Finished",.CheckBoxFinished

GroupBox 142,145,130,77,"",.GroupBox4

OptionGroup .Group3

OptionButton 150,154,80,21,"Beam 1",.OptionButtonBeam1

OptionButton 150,175,90,21,"Beam 2",.OptionButtonBeam2

OptionButton 150,196,100,21,"Both Beams",.OptionButton1

End Dialog

Dim dlg As UserDialog

dlg.TextBoxNomCtrAZ = CStr(NominalCenterAZ)

dlg.TextBoxNomCtrEL = CStr(NominalCenterEL)

dlg.TEXTBOXX = FF_XSHIFT

dlg.TEXTBOXY = FF_YSHIFT

dlg.TEXTBOXZ = FF_ZSHIFT

dlg.CheckBoxZoomIn = 0

ZoomIn = False

dlg.group1 = 0

dlg.group2 = 0

dlg.checkboxFinished = 0

Do While done = False

MEAS_PROBE_MODEL_NONE

dlg.TextBoxNomCtrAZ = CStr(NominalCenterAZ)

dlg.TextBoxNomCtrEL = CStr(NominalCenterEL)

Dialog dlg

Select Case dlg.group2

Case 0

BandSelect = 3

NominalCenterAZ = 0.323

NominalCenterEL = 1.811

Case 1

BandSelect = 6

NominalCenterAZ = 1.6867

NominalCenterEL = -1.6867

Case 2

BandSelect = 7

NominalCenterAZ = 0.974

NominalCenterEL = 0

Case 3

BandSelect = 9

NominalCenterAZ = 0

NominalCenterEL = -0.974

End Select

Select Case dlg.group3

Case 0

SELECT_BEAM 1

Case 1

SELECT_BEAM 2

Case 2

SELECT_BEAM 1

SELECT_BEAM 2, 1

End Select

If dlg.checkboxzoomin = 0 Then

ZoomIn = False

FF_HSPAN = "20 deg"

FF_VSPAN = "20 deg"

FF_HCENTER = 0

FF_VCENTER = 0

Else

ZoomIn = True

FF_HSPAN = "8 deg"

FF_VSPAN = "8 deg"

FF_HCENTER = dlg.TextBoxNomCtrAZ

FF_VCENTER = dlg.TextBoxNomCtrEL

End If

FF_PLOTCENTERING_OFF

FF_XSHIFT = CDbl(dlg.textboxx)

FF_YSHIFT = CDbl(dlg.textboxy)

FF_ZSHIFT = CDbl(dlg.textboxz)

If dlg.group1 = 0 Then

AzCut = True

FF_HCUT

Else

AzCut = False

FF_VCUT

End If

If dlg.CheckboxFinished = 1 Then

done = True

End If

Loop

ExitLoop:

CLEAR_ALL_PLOTS

ResultsArray(0) = dlg.group2

ResultsArray(1) = dlg.group3

ResultsArray(2) = NominalCenterAZ

ResultsArray(3) = NominalCenterEL

GetOptimalXYZ= ResultsArray

End Function

# Appendix C- Nominal Beam Center Coordinates