ALMA Band 9 Cartridge S/N 001

Acceptance In-House Test Report

FEND-40.02.09.00-044-B2-REP

Version: Draft 2

Status:  Draft (Draft, Pending, Approved, Released, or Obsolete)

2006-10-13

Prepared By:. Name(s) and Signature(s); Organization; Date. A. Koops B. Jackson A. Baryshev; RuG / NOVA SRON / NOVA RuG / SRON / NOVA; 2006-10-13. Approved By:. Name and Signature; Organization; Date. H. Rudolf K. Saini; ESO NROA; yyyy-mm-dd. Released By:. Name and Signature; Organization; Date. xxx; xxx; yyyy-mm-dd.

Change Record

| Version | Date | Affected Section(s) | Change Request # | Reason/Initiation/Remarks |
| --- | --- | --- | --- | --- |
| Draft 0 | 2006-04-12 | all |  | New document |
| Draft 1 | 2006-04-13 | all |  | Updated document references. |
| Draft 2 | 2006-04-18 | all |  | Output power stability results added, room-temperature health check reference, summary section, test report references. |
| Draft 3 | 2006-06-06 |  |  | Acceptance In-House Inspection Report added (cartridge mass) |
| B Draft 0 | 2006-09-21 | all |  | Updated to include cartridge test results after LO upgrade. |
| B Draft 1 | 2006-10-09 | all |  | Updated phase stability, heterodyne sensitivity, output power stability, acceptance in-house inspection test reports / results / related NCRs |
| B Draft 2 | 2006-10-13 | all |  | Added cartridge beam-pattern results, updated detailed test report list with new heterodyne sensitivity report revision, status of all reports |

Table of Contents

1	Description	4

2	Related Documents and Drawings	5

3	Configuration Item Identification and Test Result Summary	7

4	Acceptance Test Results	9

# Description

## Purpose

This document summarizes the acceptance in-house test results of a particular Band 9 Pre-Production Cartridge.

In particular, this document includes:

- an identification of the configuration status of the cartridge;

- a partial design verification matrix summarizing the acceptance in-house test results for this specific Band 9 Cartridge, including the verification method, specified and measured performance figures, and references to relevant detailed test reports; and

- a summary of the status of any NCRs and/or RFWs that are applicable to the acceptance tests of this cartridge.

Detailed results of individual tests, including plotted and tabulated results, are contained in the test reports referred to in the design verification matrix.

## Scope

This test report is specific to a particular Band 9 cartridge, the configuration of which is defined in Section 3.1 of this report.

Detailed test reports and electronic data files are contained in a single zip file attached to this report (RD 04).

Cartridge operating procedures are contained in the Band 9 Cartridge Operating Manual (RD 05), while optimum bias conditions for this particular Band 9 cartridge are contained in the appropriate Tuning Table Document (RD 10).

# Related Documents and Drawings

## Reference and Applicable Documents

| No | Document Title | Date | Reference |
| --- | --- | --- | --- |
| AD 01 | Band 9 Cartridge Technical Specifications | 2003-12-17 | FEND-40.02.09.00-002-A-SPE |
| AD 02 | Band 9 Cartridge Statement of Work | 2003-10-06 | FEND-40.02.09.00-003-B-SOW |
| AD 03 | Band 9 Cartridge Acceptance Procedure | 2004-11-05 | FEND-40.02.09.00-018-A-PRO |
| RD 04 | Band 9 Cartridge SN 001 Test Reports and Electronic Data | TBD | FEND-40.02.09.00-044-B-REP.zip |
| RD 05 | ALMA Band 9 Cartridge Operating Manual | 2006-10-09 | FEND-40.02.09.00-048-A-MAN |
| RD 06 | ALMA Band 9 Cartridge S/N 001 Configuration Overview | 2006-09-05 | S-00051-F |
| AD 07 | Band 9 Cartridge Acceptance Test Plan | 2006-01-25 | FEND-40.02.09.00-020-A-PLA |
| RD 08 | ALMA Band 9 Cartridge Test and Inspection Overview | 2006-09-18 | FEND-40.02.09.00-012-A-PLA |
| RD 09 | ALMA Band 9 Cartridge S/N 001 Compliance Matrix | 2006-06-06 update | FEND-40.02.09.00-046-A-GEN |
| RD 10 | ALMA Band 9 Cartridge S/N 001 Tuning Table Document | TBD | FEND-40.02.09.00-062-A-MAN |

## Detailed Test Reports

The following detailed test reports are contained in RD04 and are referred to by their document number (not including revision number) in the compliance matrix.

| Doc. Number | Document Title | Date | Status | Comments |
| --- | --- | --- | --- | --- |
| S-00044-G | ALMA Band 9 Cartridge S/N 001 Heterodyne Sensitivity Test Report | 2006-10-11 | Released |  |
| S-00045-E | ALMA Band 9 Cartridge S/N 001 Total Output Power Test Report | 2006-09-21 | Released |  |
| S-00046-E | ALMA Band 9 Cartridge S/N 001 Cold Health Check Test Report | 2006-09-21 | Released |  |
| S-00047-E | ALMA Band 9 Cartridge S/N 001 Output Power Stability Test Report | 2006-10-02 | Released |  |
| S-00048-C | ALMA Band 9 Cartridge S/N 001 Cross Polarization Test Report | 2006-09-20 | Released | Tested with WCA 9.1 (config. of 2006-03-24) |
| S-00049-C | ALMA Band 9 Cartridge S/N 001 Gain Compression Test Report | 2006-09-20 | Released | Tested with WCA 9.1 (config. of 2006-03-24) |
| S-00050-E | ALMA Band 9 Cartridge S/N 001 Signal Path Phase Stability Test Report | 2006-10-02 | Released |  |
| S-00052-0 | ALMA Band 9 Backpiece S/N S-05-0-012 FTS and Sideband Ratio Test Report | 2006-09-20 | Released | Mixer Pol. 0 (from RD06) |
| S-00054-0 | ALMA Band 9 Backpiece S/N S-05-0-009 FTS and Sideband Ratio Test Report | 2006-09-20 | Released | Mixer Pol. 1 (from RD06) |
| S-00056-C | ALMA Band 9 Cartridge S/N 001 Room Temperature Health Check Test Report | 2006-04-18 | Draft 0 | Tested with WCA 9.1 (config. of 2006-03-24) to be repeated w WCA9.3 |
| S-00062-0 | ALMA Band 9 Cartridge S/N 001 Cold Beam Pattern Test Report | 2006-10-13 | Draft 0 | Tested with WCA 9.1, x6 multipliers (config. of 2005-12-02) |
| S-00068-A | ALMA Band 9 Cartridge S/N 001 Acceptance In-House Inspection Report | 2006-10-02 | Draft 1 | Mass determined for config. of 2006-07-19 Inspection to be repeated |
| S-00116-0 | ALMA Band 9 Cartridge Baseplate S/N 001 Vacuum Leak-Rate Test Report | 2006-09-20 | Released | Tested with config. of 2006-07-19 (48 µm mica windows) |

## Abbreviations and Acronyms

AD	Applicable Document

ESO	European Southern Observatory

NCR	Non-Conformance Report

NOVA	The Netherlands Research School for Astronomy

RD	Reference Document

RFW	Request for Waiver

RuG	University of Groningen

SRON	SRON Netherlands Institute for Space Research

# Configuration Item Identification and Test Result Summary

## Configuration Item Identification

| Product Name | Band 9 Cartridge |
| --- | --- |
| Serial Number | 001 |
| Production Report Reference | RD06 |

## Applicable NCR’s and RFW’s

| NCR Number | RFW Number | NCR Name | Non-Conformance | Report Reference | Major or Minor | Status |
| --- | --- | --- | --- | --- | --- | --- |
| B9-NCR0044 | FEND-40.02.09.00-055-A-RFW | Sideband mismatch | sideband mismatch > 3 dB for some combinations of FIF > 9.2 GHz and FLO < 634 GHz | Section 4.1, Line 5.2 S-00052 S-00054 | Minor | pending |
| B9-NCR0045 | FEND-40.02.09.00-057-A-RFW | IF power variations | power variation out of spec due to excess gain slope originally observed gain slope partially corrected with slope-correcting amplifiers | Section 4.1, Line 5.3.2 S-00044 | Major | pending |
| B9-NCR0046 | FEND-40.02.09.00-052-A-RFW | Cross polarization | measured cross-polarization level ~ -17-18 dB | Section 4.1, Line 5.7.4 S-00048 | Major | pending |
| B9-NCR0049 | FEND-40.02.09.00-053-A-RFW | Cartridge mass | Added cooled mass ~ 2.05 kg vs. spec = 2.0 kg | S-00068 | Minor | pending |
| B9-NCR0067 | FEND-40.02.09.00-063-A-RFW | Vacuum leak-rate | Measured vacuum leak-rate exceeds specified value | S-00116 | Major | open |

## Verification Conclusions

| Product Compliant: | RFWs pending |
| --- | --- |

## Comments

In-house acceptance tests of cartridge #1 were completed using WCA 9 S/N 001 and x6 LO multipliers. A number of these tests were affected by the performance of the prototype local oscillator in WCA 9 S/N 001, including excess LO noise and drifts in the LO output power. An upgrade to the LO configuration has resulted in partial retesting using WCA 9 S/N 003. Only tests directly influenced by the LO performance have been repeated. The acceptance in-house inspection and room-temperature health check will be repeated/completed prior to delivery.

# Acceptance Test Results

The following partial design verification matrix summarizes the results of acceptance testing of the Band 9 Cartridge, including the requirements, the measured performance, the verification method, the level at which that verification took place (i.e. sub-assembly and/or cartridge), and a reference to the detailed test report in which relevant additional information can be found.  These test reports are identified in Section 2.2.

Only those requirements that have been verified by test, or which are identified in the Acceptance Test Plan (AD07) as requiring verification by test, are included here.  A full summary of the compliance status of the cartridge, including compliance with relevant applicable documents to the Cartridge Technical Specifications (AD01), including ICDs, is included in RD09.

With respect to the verification method, the following abbreviations are used:

- R	Review of Design

- I	Inspection

- A	Analysis

- T	Test

- Tsup	Supplementary Verification by Test (at a lower level of integration, or as input to an analysis)

- S	Similarity (reference to a previous project)

- Sn	Similarity (reference to a previous cartridge, where n is the serial number of that cartridge)

- N/A	Not applicable

Review of design shall take place at the PDR and/or CDR.

NCR’s or RFW’s that result from non-compliances in this verification are included in Section 3.2.

## Summary of Test Results

| Sec. in AD01 | Name | ALMA Specification # | Req. | Result | Compliant | Verification Method / Level |  | Test Method | Detailed Report | Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  | Sub-Ass. | Cart. |  |  |  |
| 3 | General Requirements |  | -- | -- | -- | -- | -- | -- | -- |  |
| 3.1 to 3.5 | Operations mode to Metric Hardware | FEND-40.02.09.00-00010-00/I to FEND-40.02.09.00-00110-00/R | various | not tested | -- | -- | -- | -- | -- | acceptance test not required in AD07 |
| 4 | Functional requirements |  | -- | -- | -- | -- | -- | -- | -- |  |
| 4.1 to 4.2 | Mixer type to Mixing scheme | FEND-40.02.09.00-00120-00/R to FEND-40.02.09.00-00130-00/R | various | -- | C | -- | R | -- | -- |  |
| 4.3 | Frequency coverage |  | -- | -- | -- | -- | -- | -- | -- |  |
| 4.3.1 | RF input port | FEND-40.02.09.00-00140-00/R | 602-720 GHz | -- | C | -- | I | -- | S-00046 S-00044 |  |
| 4.3.2 | LO input port | FEND-40.02.09.00-00150-00/R | 614-708 GHz | -- | C | -- | I | -- | S-00046 S-00044 |  |
| 4.3.3 | IF output bandwidth, Fcentre | FEND-40.02.09.00-00160-00/R | 4-12 GHz | -- | C | -- | I | -- | S-00044 |  |
| 4.4 | Polarization State | FEND-40.02.09.00-00170-00/R | 2 orthogonal, linear polarizations | -- | C | -- | I | -- | S-00048 |  |
| 5 | Performance requirements |  | -- | -- | -- | -- | -- | -- | -- |  |
| 5.1 | Cartridge noise performance | FEND-40.02.09.00-00180-00/T | TN,rec < 168 K over 80% of band TN,rec < 250 K over full band | Total Power: TN < 168 K for FLO = 630...710 GHz TN < 250 K for FLO = 614, 622 GHz Variation with IF: TN < 1.25TN,req (210 and 312 K) for all FIF, FLO | C | -- | T | TN,rec | S-00044 | measured in 50 MHz bandwidth |
| 5.2 | Sideband mismatch | FEND-40.02.09.00-00190-00/T | < 3dB across IF band | Pol. 0 and 1: < 3.8 dB for all FLO, FIF < 3 dB for all FIF for FLO > 634 GHz < 3 dB for all FLO for FIF < 9.2 GHz | PC | Tsup | A | analysis of mixer FTS | Pol. 0 S-00052 Pol. 1 S-00054 |  |
| 5.3 | Cartridge IF output |  | -- | -- | -- | -- | -- | -- | -- |  |
| 5.3.1 | IF output power | FEND-40.02.09.00-00200-00/T | -- | -- | -- | -- | -- | -- | -- |  |
|  | Total in-band power |  | -37 to -24 dBm in 4-12 GHz | Pout,hot,4-12 = -30.5 to -27 dBm at all FLO, for both polarizations | C | -- | T | TN,rec | S-00045 | bandwidth = 3.7 - 12.6 GHz |
|  | Total power |  | < -22 dBm in 10 MHz to 18 GHz | Pout,hot,0-18 = -31 to -27.5 dBm at all FLO, for both polarizations | C | -- | T | TN,rec | S-00045 | bandwidth = 0 - 18 GHz |
| 5.3.2 | IF power variations | FEND-40.02.09.00-00210-00/T | -- | -- | -- | -- | -- | -- | -- |  |
|  | 2 GHz sub-band |  | +/- 2dB | Pol. 0: 3.2-4.1 dB p-p Pol. 1: 4.3-4.8 dB p-p | NC | -- | T | TN,rec | S-00044 | measured in 50 MHz bandwidth; variation within 2GHz compromised due to gain slope |
|  | Full IF band |  | +/- 3dB | Pol. 0: 3.7-5.6 dB p-p Pol. 1: 6.4-7.8 dB p-p | NC | -- | T | TN,rec | S-00044 | measured in 50 MHz bandwidth; gain slope due to SIS mixer impedance mismatch and isolator loss |
| 5.4 | Gain saturation | FEND-40.02.09.00-00230-00/T | < 3% for 373 K vs. 77K | measured compression = -1% to 3% | C | -- | T | TN,rec + external signal source | S-00049 | negative compression implies increase in response at 373K |
| 5.5 | Amplitude stability | FEND-40.02.09.00-00240-00/T | σ2 < 4x10-7 | all FLO, both polarizations σ2 < 3x10-7 | C | -- | T | Allan variance | S-00047 |  |
| 5.6 | Phase stability | FEND-40.02.09.00-00250-00/T | < 2.5deg up to 300 sec. (16 msec samples) | 1.2-2.4º peak-to-peak after removing 0.8-1.4º variation at 1 Hz | C | -- | T | signal path phase sampled over time | S-00050 | sampled at 300 Hz with dead-time every 20 samples p-to-p variation includes 0.8-1.4º at 1 Hz (cryo-cooler) |
| 5.7 | Optics |  | -- | -- | -- | -- | -- | -- | -- |  |
| 5.7.1 | RF beam efficiency | FEND-40.02.09.00-00260-00/T | coupling to subreflector > 90% | > 96% | C | -- | T | near-field beam measurement | S-00062 |  |
| 5.7.2 | Polarization alignment | FEND-40.02.09.00-00270-00/T | E-vector of polarization 0 towards centre of cryostat | as specified | C | -- | I | TN,rec + rotating grid | S-00048 |  |
| 5.7.3 | Polarization align. accuracy | FEND-40.02.09.00-00280-00/T | < 2 deg | < 1.3 deg | C | -- | T | TN,rec + rotating grid | S-00048 |  |
| 5.7.4 | Cross-polarization | FEND-40.02.09.00-00290-00/T | < -20dB | -16 to -18.5 dB | NC | -- | T | TN,rec + rotating grid | S-00048 | see comments in S-0048 and related RFW |
| 5.8 | Stabilization time | FEND-40.02.09.00-00300-00/T to FEND-40.02.09.00-00305-00/T | various | not tested | -- | -- | -- | -- | -- | acceptance test not required in AD07 |
| 6 | Mechanical and Electrical Requirements |  | -- | -- | -- | -- | -- | -- | -- |  |
| 6.1 | Mass | FEND-40.02.09.00-00310-00/T | < 2 kg | added mass ~ 2.05 kg | NC | -- | I | mechanical measurement | S-00068 | cooled mass added to blank cartridge body |
| 6.2 | Eigenfrequency | FEND-40.02.09.00-00320-00/A | > 70 Hz | not tested | -- | -- | -- | -- | -- | test to be performed on later cartridge to be reviewed at CDR |
| 6.3 | Volume | FEND-40.02.09.00-00330-00/I | within D = 170 mm and H = 475 mm | as specified | C | -- | I | visual inspection | S-00064 |  |
| 6.4 to 6.7 | Orientation to Connectors and RF ports | FEND-40.02.09.00-00340-00/T to FEND-40.02.09.00-00400-00/R,I | various | not tested | -- | -- | -- | -- | -- | acceptance test not required in AD07 |
| 7 | Operating conditions |  | -- | -- | -- | -- | -- |  | -- |  |
| 7.1 | Local oscillator | FEND-40.02.09.00-00420-00/T to FEND-40.02.09.00-00450-00/T | various | not tested | -- | -- | -- | -- | -- | acceptance test not required in AD07 |
| 7.2 | Thermal environment | FEND-40.02.09.00-00460-00/T | -- | -- | -- | -- | -- | -- | -- |  |
|  | temperature ranges |  | < 4K 10-18K 80-130K | cartridge tested at: T4K, cryostat = 4.0 K T15K, cryostat ~ 12 K T110K, cryostat ~ 67 K T4K, cartridge = 4.1 K T15K, cartridge = 12-13 K T110K, cartridge = 66-69 K | PC | -- | R, Tsup | various | S-00044 S-00045 S-00046 S-00047 | Cartridge tested at a single temperature (T4K, cryostat = 4 K, all other temperatures as defined by the test cryostat). |
|  | temperature stability |  | 4K +/- 2 mK 15 K +/- 50 mK 110K +/- 100 mK | -- | -- | -- | -- | -- | -- | Cartridge output power, phase stability measured in test cryostat. Cartridge test cryostat not representative of ALMA cryostat |
|  | warm up during operation |  | shall survive warmup to 300K | not tested | -- | -- | -- | -- | -- | verification not required in AD07 |
| 7.3 | Vacuum conditions | FEND-40.02.09.00-00470-00/R | Component outgassing shall comply with Dewar ICD | helium leak-rate after 20 min. exposure to 1 atm. helium = 1.8x10-6 mbar·L/s | NC | -- | I | -- | S-00116 | Mica windows are leak tight. Leak-rate dominated by diffusion through Viton o-rings. |
| 7.4 | Environmental operating conditions |  | -- | -- | -- | -- | -- | -- | -- |  |
| 7.4.1 | Vibration | FEND-40.02.09.00-00480-00/T | Cartridge shall survive vibration levels defined in Environmental Requirements | not tested | -- | -- | -- | -- | -- | to be tested on a later cartridge |
| 7.4.2 | Acceleration | FEND-40.02.09.00-00490-00/T | shock loading 3g (vertical) 2g (hor.) | not tested | -- | -- | -- | -- | -- | to be tested on a later cartridge |
| 7.5 | Storage/shipping conditions | FEND-40.02.09.00-00500-00/R | Cartridge shall comply with Environmental Requirements | not tested | -- | -- | -- | -- | -- | acceptance test not required in AD07 |
| 7.6 | EMC | FEND-40.02.09.00-00510-00/T | Cartridge shall comply with EMC Requirements | not tested | -- | -- | -- | -- | -- | acceptance test at PAI not required in AD07 |
| 7.7 | Monitoring and control | FEND-40.02.09.00-00520-00/R to FEND-40.02.09.00-00550-00/R | various | not tested | -- | -- | -- |  | -- | acceptance test not required in AD07 excess noise observed in readout of SIS bias current transients observed due to switching of temperature sensors heater rise/fall time much shorter than specified |
| 8 | RAMS requirements | FEND-40.02.09.00-00560-00/R to FEND-40.02.09.00-00580-00/R | various | not tested | -- | -- | -- | -- | -- | acceptance test not required in AD07 |
|  |  |  | -- | -- | -- | -- | -- | -- | -- |  |
| Non-Cartridge Specifications Identified in AD07 |  |  | -- | -- | -- | -- | -- | -- | -- |  |
|  | Cartridge cold plate alignment | FEND-40.03.00.00-00230-00/T to FEND-40.03.00.00-00235-00/T | ??? | not tested | -- | -- | -- | -- | -- | acceptance test not required in AD07 |
|  | LO phase stability | FEND-40.00.00.00-00600-00/T | ??? | not tested | -- | -- | -- | -- | -- | acceptance test not required in AD07 |
|  | LO frequency switching | FEND-40.00.00.00-00530-00/T to FEND-40.00.00.00-00590-00/T | ??? | not tested | -- | -- | -- | -- | -- | acceptance test not required in AD07 |
|  | LO phase switching | FEND-40.00.00.00-00630-00/T | ??? | not tested | -- | -- | -- | -- | -- | acceptance test not required in AD07 |